section 11 5311.10

glovebox FABRICATION

*************************************************************************************************************

Edit this template for each project.  In doing so, specifier must add job-specific requirements.  Brackets are used in the text to indicate designer choices or locations where text must be supplied by the designer.  Once the choice is made or text supplied, remove the brackets.  The specifications must also be edited to delete specification requirements for processes, items, or designs that are not included in the project -- and specifier’s notes such as these.  To seek a variance from requirements in the specifications that are applicable, contact the Engineering Standards Manual Mechanical POC. Please contact POC with suggestions for improvement as well.

This specification was prepared by an organization operating under a quality assurance program that meets the requirements of DOE Order 414.1D, ASME NQA-1, and LANL SD330 (suitable for ML-1 through ML-4 projects).  Implementation of this specification requires modification to the specification to meet project-specific requirements.  Responsibility for application of this specification to meet project-specific requirements lies with the organization modifying or implementing the specification.  The organization modifying the specification shall apply a graded approach to quality assurance based on the ML designation of the project.  When this specification is used with nuclear facilities subject to DOE Order 414.1D, modification to this specification must be performed by an individual or organization operating under a quality assurance program that meets the requirements of same.

This specification is a general specification covering a wide range of nuclear materials glovebox applications.  It is to be used for fabrication of gloveboxes when combined with build-to-print fabrication drawings.  If a design/build procurement is needed, this section should be edited in conjunction with Section 11 5311.08 – Glovebox Design.

This specification primarily defines requirements for fabrication of stainless steel gloveboxes used for confinement of nuclear materials.  Other materials may be used for the fabrication of the glovebox, including aluminum, provided that the process and corrosive requirements inside the glovebox are met.  Redefine the specification requirements when specifying a glovebox fabricated from a material different than stainless steel.  If the process requires coating, refer to Section 11 5311.19 – Chemical Resistant Coatings for Gloveboxes.

This specification is most applicable to new acquisitions, but may have application with modifications or repair work to existing gloveboxes.

*************************************************************************************************************

general

section includes

Glovebox Shells

Open-Front Boxes

Airlocks

Glovebox Support Stands

Gloveports

Windows

Filter Housings

Shell Penetrations

Access and Service Panels

Material Transfer Devices

Radiation Shielding

Glovebox Linings

Pressure Relief Devices

Other Glovebox Appurtenances

scope

This section establishes the technical requirements for the materials of construction, fabrication, testing, shipment, and quality assurance (QA) of gloveboxes (GBs), their support stands, and components or appurtenances.

The technical requirements of this specification are applicable to GBs used for the primary and secondary confinement of nuclear materials.  Additional technical requirements are provided in the Drawings.  The following is a summary of supplier responsibilities described in this specification:

Fabricate GBs in accordance with this specification and the Drawings.

Procure equipment, materials, or supplies to complete the work, unless otherwise stated.

Test and inspect as required by this specification.

Furnish the data required by this specification to document that required tests and inspections have been performed.

Package, ship, and deliver GBs.

Provide LANL full access to the facility for performing random or scheduled inspections and/or surveillance of work performed.

Provide LANL with a lower tier services plan including the name, address, telephone number, and point of contact for outside services that the supplier intends to use on this project.  Identify the specific work requirements of this specification that will be performed by those outside services.

Provide a fabrication schedule showing fabrication steps, hold points, tests, and inspections.  Provide a revised fabrication schedule after any modification to the Subcontract document, which revises the required delivery date, or when other approved LANL changes otherwise change a scheduled assembly step hold point, test, or inspections.

Provide seven (7) working days advance notice of a hold point activity requiring LANL witness or inspection.

LANL GB Detail Drawings

Attachment 1 of this Section provides a list of LANL GB detail drawings often referenced by the Subcontract Drawings provided to the supplier for fabrication of GBs.

*************************************************************************************************************

Edit applicable related sections defined below to meet the project requirements associated with the GB to be fabricated.

*************************************************************************************************************

related sections

Section 01 1116: Work by Owner

Section 01 4444: Offsite Welding and Joining Requirements

Section 01 6000: Product Requirements

Section 01 7839: Project Record Documents

Section 11 5311.19: Chemical Resistant Coatings for Gloveboxes

definitions

Access Panel: A removable and resealable panel used for interior access.

Airlock: A transition enclosure for material movement into and out of the GB that maintains the primary confinement.  The term “transfer airlock” is sometimes used interchangeably with airlock.  An airlock is sometimes purged with inert gasses.

Certificate of Conformance:  Document certified by a competent authority that the supplied good or service meets the required specifications.  “Certificate of Compliance” is similar but attests to the supplier of goods or services meeting specification.

Certified Material Test Reports (CMTR), Mill Traceability Report (composition), and Material Test Report (mechanical properties).  All must have actual test results.  CMTRs shall be certified by a nationally accredited lab (notary not required.

Drawings: The drawings of the Subcontract.

Glovebox (GB):  A controlled environment enclosure providing primary confinement from the work area.  Operations inside GBs are performed through sealed glove openings for the protection of the worker, the environment, and/or the process.

Hood: An enclosure similar to a chemical fume hood.  A non-isolated enclosure for controlled access to a GB that may also be used independently as a stand-alone unit or in a line with other hoods.  Confinement is achieved through airflow in a hood.  Hoods can also be used for low-level analytical chemistry operations.  Can also be referred to as an Open-Front GB, Introductory GB, or Radio-Benches.  For the purposes of this specification only, the term “hood” will be used interchangeably with Open-Front GB, Introductory GB, and Radio-Bench.

Primary Confinement: The barrier (structure) that is directly in contact with bulk radioactive material.  The barrier, that if breached exposes the bulk radioactive material.  Primary confinement is pipes and vessels in a tritium system.  Primary confinement in an SNM process may be process piping and vessels or may be the GB when solids are exposed in machining and handling processes.

Secondary Confinement: Secondary confinement is a structure erected around primary confinement for the purposes of creating a barrier to block migration of unanticipated and anticipated breaches of the primary confinement.  GBs are secondary confinement in a tritium system or a wet chemistry SNM process where they isolate releases of radioactive materials when the process piping must be opened.

Shielded GBs: A shielded GB is a GB provided with radiation shielding.  Depending on process conditions, certain GBs may be shielded by the addition of gamma shielding and/or neutron shielding covering the front, sides, back, and bottom of the GB as required.  This shielding may be covered by stainless steel.

Supplier: A Subcontractor providing services to LANL.

Acronyms

AGS: American Glovebox Society.

AISI: American Iron and Steel Institute

ANSI: American National Standards Institute

ASME: American Society of Mechanical Engineers

ASNT: American Society for Nondestructive Testing

ASTM: ASTM International (formerly American Society for Testing and Materials)

AWS: American Welding Society

B&PVC: ASME Boiler and Pressure Vessel Code

CD: capacitive discharge

CFR: Code of Federal Regulation

CJP: complete joint penetration

CMTR: Certified Material Test Report

CoC: Certificate of Conformance

CRL: Central Research Laboratories

DOE: Department of Energy

GB: glovebox

LANL: Los Alamos National Laboratory

NFPA: National Fire Protection Association

NDE: nondestructive examination

NQA: Nuclear Quality Assurance

PJP: partial joint penetration

PQR: (Welding) Procedure Qualification Record

QA: quality assurance

QC: quality control

RT: radiographic testing

SAE: Society of Automotive Engineers

SDDR: Supplier Deviation Disposition Request

SNM: Special Nuclear Material

SSPC: Steel Structures Painting Council

UL: Underwriters Laboratories

WPS: Welding Procedure Specification

WPQR: Welder Performance Qualification Record

Regulatory and Technical requirements

Codes, specifications, and standards referred to by number or title form a part of this specification to the extent required by the following references and others that may exist in this document.  Use codes, specifications, and standards referenced below of the latest revision at the time of award of Subcontract, unless otherwise stated below.

10 CFR 830.122: Quality Assurance Criteria

AGS-G006-2005: Standard of Practice for the Design and Fabrication of Nuclear-Application Gloveboxes

AGS-G004-2014: Standard of Practice for Leak Test Methodologies for Gloveboxes and other Enclosures

ASME AG-1-2015: Code on Nuclear Air and Gas Treatment

ASME B18.2.1: Square, Hex, Heavy Hex, and Askew Head Bolts Hex, Heavy Hex, Hex Flange, Lobed Head and Lag Screws (Inch Series)

ASME Boiler and Pressure Vessel Code (B&PVC), Section II – Part C: Materials Specifications – for Welding Rods, Electrodes, and Filler Metals

ASME B&PVC, Section V: Nondestructive Examination

ASME B&PVC, Section VIII – Division I, Rules for Construction of Pressure Vessels

ASME B&PVC, Section IX: Qualification Standard for Welding and Brazing Procedures, Welders, Brazers, and Welding and Brazing Operators

ASME NQA-1: Quality Assurance Requirements for Nuclear Facility Applications

ASNT SNT-TC-1A: Recommended Practice for Personal Qualification and Certification in Nondestructive Testing

ASTM A36: Standard Specification for Carbon Structural Steel

ASTM A182: Standard Specification for Forged or Rolled Alloy Steel Pipe Flanges, Forged Fittings and Valves, and Parts for High-Temperature Service

ASTM A193: Standard Specification for Alloy-Steel and Stainless Steel Bolting for High-Temperature Service or High Pressure Service and Other Special Purpose Applications

ASTM A194: Standard Specification for Alloy Steel and Stainless Steel Bolting for High- Temperature or High- Pressure Service and Other Special Purpose Applications

ASTM A240: Standard Specification for Chromium and Chromium-Nickel Stainless Steel Plate, Sheet and Strip for Pressure Vessels  and for General Applications

ASTM A276: Standard Specification for Stainless Steel Bars and Shapes

ASTM A307: Standard Specification for Carbon Steel Bolts and Studs, 60,000 PSI Tensile Strength

ASTM A325: Standard Specification for Structural Bolts, Steel, Heat Treated, 120/105 ksi Minimum Tensile Strength

ASTM A354: Standard Specification for Quenched and Tempered Alloy-Steel Bolts, Studs, and Other Externally Threaded Fasteners

ASTM A479: Stainless Steel Bars and Shapes for use in Boilers and other Pressure Vessels

ASTM A480: General Requirements for Flat Rolled Stainless and Heat-Resisting Steel Plate, Sheet, and Strip

ASTM A511: Standard Specification for Seamless Stainless Steel Mechanical Tubing

ASTM A554: Standard Specification for Welded Stainless Steel Mechanical Tubing

ASTM A563A: Standard Specification for Carbon and Alloy Steel Nuts

ASTM A572: Standard Specification for High-Strength Low-Alloy Columbium-Vanadium Structural Steel

ASTM A588 Standard Specification for High-Strength Low-Alloy Structural Steel, up to 50 ksi [345 MPa] Minimum Yield Point, with Atmospheric Corrosion Resistance

ASTM A1008: Standard Specification for Steel, Sheet, Cold-Rolled, Carbon, Structural, High-Strength Low-Alloy, High-Strength Low-Alloy with Improved Formability, Solution Hardened, and Bake Hardenable

ASTM B209: Standard Specification for Aluminum and Aluminum-Alloy Sheet and Plate

ASTM B749: Standard Specification for Lead and Lead Alloy Strip, Sheet, and Plate Products

ASTM C1036: Standard Specification for Flat Glass

ASTM C1172: Standard Specification for Laminated Architectural Flat Glass

ASTM E498: Standard Practice for Leaks Using the Mass Spectrometer Leak Detector or Residual Gas Analyzer in the Tracer Probe Mode

ASTM E499: Standard Practice for Leaks Using the Mass Spectrometer Leak Detector in the Detector Probe Mode

ASTM F593: Standard Specification for Stainless Steel Bolts, Hex Cap Screws, and Studs

ASTM F594: Standard Specification for Stainless Steel Nuts

AWS A2.4: Standard Symbols for Welding, Brazing and Nondestructive Examination

AWS A5.32:  Welding Consumables – Gases and Gas Mixtures for Fusion Welding and Allied Processes

AWS B2.1: Specification for Welding Procedure and Performance Qualification

AWS D1.1: Structural Welding Code Steel

AWS D1.6:  Structural Welding Code – Stainless Steel

DOE Order 414.1D Quality Assurance

DOE-STD-1066 Fire Protection

NFPA 801: Fire Protection for Facilities Handling Radioactive Materials

SAE J429: Mechanical and Material Requirements for Externally Threaded Fasteners

SSPC SP-1: Solvent Cleaning

SSPC SP-6: Commercial Blast Cleaning

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author is required to define many aspects of the GB design.  Designate management level of the GB and components in accordance with AP-341-502, Management Level Determination.  Following the designation of management level, interpret the requirements for the designated management level and define those requirements in this specification.

NOTE: This section was also written to minimize the need for bundling with other sections.  As an example, the following article repeats material found in Section 01 3300 on submittals (although 01 3300 is referenced later in this Section).  Thus, if 01 3300 is included in the procurement, articles such as the following must be reviewed against it to ensure alignment (or streamlined).

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

submittals

Include LANL Subcontract Number, Glovebox Title, and Drawing Number on correspondence.

Provide submittals listed in Attachment 2 to LANL for review before shipment of the GB.  The Subcontractor shall allow a minimum of [14] calendar days for return of submittals.

### LANL will utilize the following "Status Codes" to indicate the status of submittals resulting from the review, and the action required of the Subcontractor.

- Status Code 1:	Work may proceed.

- Status Code 2:	Revise and resubmit.  Work may proceed subject to resolution of indicated comments.

- Status Code 3:	Revise and Resubmit.  Work May Not Proceed.

- Status Code 4:	Review not required. Work may proceed.

    - Use a routing sheet with all resubmittals indicating each item's submittal number and type suffixed "R1" for the first resubmittal, "R2" for the second resubmittal, and so forth.

### Do not fabricate products or begin Work that requires submittals before such submittals are approved.

#### Exceptions: Field tests, inspection reports, and Subcontract closeout submittals.

Submittal for SUBContract Closeout

    - Provide binders with clear pockets on front and binding edge for insertion of titles. Binders shall be not more than 75 percent full.

#### Internally subdivide the binder contents with permanent page dividers, logically organized as described below, with tab titling clearly printed under reinforced laminated plastic tabs.

Prepare a Table of Contents for each volume, with each product or system description identified, printed on 24-pound white paper.

Part 1: Directory-listing names, addresses, and telephone numbers of sub-tier Subcontractors and major equipment suppliers.

#### Part 2: Operation and maintenance instructions arranged by specification section.  For each category, identify names, addresses, and telephone numbers of sub-tier Subcontractors and suppliers.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine whether the kick-off meeting is required and the appropriate time for the kick-off meeting.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

kick-off meeting

    - Project Engineer will hold a half-day kick-off meeting at Los Alamos, NM within 10 working days after award of the Subcontract.  The kick-off meeting may be waived at the discretion of LANL.  Provide technical documentation required for submittal at the kick-off meeting conference even if a kick-off meeting is not held.  Review the following at the kick-off meeting:

Subcontract document provisions.

Technical specifications.

Drawings.

Fabrication schedule.

Supplier’s lower tier services plan.

Supplier’s QA manual.

Required procedures.

Welding and NDE Personnel List.

Philosophy of shop traveler.

Identify critical submittals that require LANL approval prior to start of procurement and/or fabrication.  (i.e., WPS, welder qualifications/continuity, filler materials, etc.)

Shop traveler, customized by the Fabricator for the project.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine and specify schedule requirements including required delivery date if critical.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Schedule

    - Provide a fabrication schedule showing fabrication steps, hold points, tests, and inspections.  Provide fabrication schedule to LANL, for approval, at the kick-off meeting.

Provide a revised fabrication schedule for LANL approval within seven (7) working days of a modification to the Subcontract document, which revises the required delivery date, or when other approved LANL changes otherwise change a schedule assembly step hold point, test, or inspection.

Provide seven (7) working days advance notice of a hold point activity requiring LANL witness or inspection.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine whether a shop traveler is required for submittal with the fabrication of the GB.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Shop Traveler

    - Use a shop traveler system and the Drawings to transform the technical requirements of this specification into specific work instructions, which indicate the fabrication and inspection sequence and identify hold points.

Provide a shop traveler procedure.  Include in the procedure a description of the preparation, use, and monitoring of the shop traveler.  Also include in-process tracking of items, processes, hold points, and inspections.  Submit for approval at the kickoff meeting, the shop traveler, customized for the subject project.  Provide the shop traveler form as an attachment to the shop traveler procedure.

After work is finished, submit the completed shop traveler as part of the QA Document Package.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine and specify any additional fabrication hold points required on a project-by-project basis.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Fabrication Hold Points

Hold points are required during the fabrication process to allow inspection, verification, or approval by LANL before the supplier does further work.  Identify hold points on the shop traveler and make provision for LANL signoff.  The minimum required hold points are:

Weld Fit-up and dimensional inspection – to inspect all joint fit up for correct gap and joint preparation, 5% of weld tacks, correct orientation of windows/ports/appurtenances, and overall dimensions. The Design Agency will identify critical dimensions for this inspection. Overall dimensions to be verified using the Drawings.

Note: Any inspection deficiencies/discrepancies must be resolved before fabrication proceeds.

Finish welds in the GB shell must be inspected before to ensure adequate size, that welds are free of undercutting and any un-fused overlap of the weld deposit.

Note: Any inspection deficiencies/discrepancies must be resolved before fabrication proceeds.

Lead Fit-up Inspection (Lead-shielded GBs only).

Window Flatness Inspection: Planar flatness to be done at postweld/pre-coating, at final acceptance, and at delivery.

Coated GBs: Shop acceptance tests for enclosures and components include Dry Film Thickness (DFT) Test, Spark Test, and Bond Test in accordance with Section 11 5311.19, Chemical Resistant Coatings.

Pre-Coating Inspection: Resolution of any non-conforming items from preliminary inspection.  Weld inspection to include sampling plan for weld studs, radiography for full penetration welds, and acceptable polishing of welds. The inspection of interior parts that will be permanently attached to the interior of the shell.

Post-Coating Inspection: Inspection of critical dimensions, overall dimensions, disposition of out of tolerance dimensions, and acceptance/rejection of dimensional control map.

GB Leak Test (Global Leak test and Helium leak test at final acceptance): Global leak test includes full volume of glovebox and any appurtenances, Helium Leak test to be performed on every weld joint (refer to a Weld Map for this test), and every attachment to the GB shell, to include Airlocks and other appurtenances.

Shop Acceptance Testing and Inspections (Attachment 4) - Inspection of enclosure welds, surface finish, and final dimensional inspection including 10% RT inspection of butt welds, weld joint verification and weld map verification, and all Leak Testing.

For hold points requiring witness or inspection at the supplier’s facility, provide seven (7) working days advance written notification to LANL so that a LANL representative may be present at the supplier’s shop to witness the activity.

Inspection Data Sheet:  The LANL representative is required to log an inspection report for each activity.  The Inspection Data Sheet is provided as Attachment 4.  The LANL representative will document inspection conformances and/or exceptions. The Project Manager will give approval to proceed with fabrication, based on inspection acceptance or rejection.

Submit the completed Inspection Data Sheet form as part of the QA Document Package.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The following quality assurance requirements are consistent with fabrication of an ML-1 or ML-2 GB.  Where other quality assurance and quality control requirements are needed, modify the following section accordingly.  For instance, the supplier may also apply a QA program in accordance with basic requirements of 10 CFR 830.122.  Add requirements for QA Programs compliant with 10 CFR 830.122 to the specification as necessary.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

quality assurance/quality control

The supplier shall implement and maintain a quality assurance program in accordance with 10 CFR 830 Quality Assurance; DOE Order 414.1D, Quality Assurance; and  Part I of ASME NQA-1-2008 with ASME NQA-1A-2009 addenda, QA Requirements for Nuclear Facility Applications, and the following additional requirements:

Suspect/counterfeit items controls and processes shall be in accordance with DOE Order 414.1D, Attachment 3, Suspect/Counterfeit Items Prevention.

GBs shall be packaged, shipped, and stored in accordance with Level B requirements of ASME NQA-1, Part II, Subpart 2.2, Quality Assurance Requirements for Packaging, Shipping, Receiving, Storage, and Handling of Items for Nuclear Power Plants.

Safety Software as defined in DOE Order 414.1D shall be in accordance with Attachment 4 (Safety Software Quality Assurance Requirements for Nuclear Facilities) of DOE Order 414.1D, Quality Assurance and ASME NQA-1, Requirements 3, and 11 and Part II, Subpart 2.7, Quality Assurance Requirements for Computer Software for Nuclear Facility Applications.  See also LANL ESM Ch. 21 Software.  Safety Software submittals shall be in accordance with Exhibit H of the Subcontract.

Commercial Grade Dedication shall be performed in accordance with ASME NQA-1, Part II, Subpart 2.14, Quality Assurance Requirements for Commercial Grade Items and Services.  Prior to fabrication, the supplier shall submit for concurrence dedication plans or procedures including the essential elements of the dedication process. Prior to shipment, the supplier shall submit other dedication records for concurrence including the following:

Commercial grade item or service procurement documents,

Technical evaluations

Critical characteristics identification and acceptance criteria

Test reports or results, inspection reports, analyses reports

Commercial grade survey reports

Source verification reports

Historical performance information

Dedication report containing sufficient data to accept the item or service.

QA Document Package: Submit documents identified in this specification as a part of the QA Document Package.  Complete three bound or stapled document packages containing these documents required “with shipment”.  Mail one package to LANL and provide the other two packages with the GB shipping crate.

MATERIAL CONTROL

Material Control Procedure

Submit to LANL, for approval, a material control procedure to be used in the execution of the work.  Describe the control methods and traceability documentation in the procedure used by the supplier to handle and monitor the use of controlled materials, such as stainless steel and welding filler rod.

Address procurement controls, segregation of materials, and traceability of materials from receipt at the shop through processing and final assembly in the procedure.  Submit this procedure to LANL for engineering review, comment, and approval, prior to procurement.

Heat Numbers: Note heat numbers on weld maps. Using a low-chloride content marking pen, identify the heat number on each piece part, and transfer the material identifications to the as-built drawings as described in § 2.3.

EXCEPTIONS, DEVIATIONS, AND CONFLICTS

Submit a written request to LANL for any proposed technical changes, exceptions, and/or deviations to this specification or the Drawings.  Submit proposed changes that affect cost or schedule in accordance with the provisions of the Subcontract document.

Supplier Deviation Disposition Request (SDDR): Provide proposed change information using the SDDR Form 2178.  Consecutively number SDDRs and submit by email for expediency with a record copy to follow by mail.  Submit one (1) copy the LANL Subcontract Administrator and a second to the LANL Subcontract Technical Representative (STR).  Do not implement proposed changes, exceptions, or deviations until the LANL STR provides written approval by means of the SDDR form.

Conflicts: Notify LANL in the event of conflicts amongst the specifications, Drawings, and/or the manufacturer’s recommended processes or instructions.  Provide notification of a conflict immediately following its discovery.  Provide notification in written form, or via phone call followed by email.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Evaluate and determine whether proposed substitutions for “or equal” products are adequate for the intended purpose.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Substitutions

Request approval for substitutions by submitting an SDDR form, where this specification or the Drawings contain a “brand name or equal” description.  Submit one (1) copy to the LANL STR.

If the supplier proposes to provide an equal to the listed brand names, the evaluation and the determination as to the quality of the proposed product will be the decision of LANL and will be based on information furnished by the supplier.

There will be no responsibility by LANL for locating or securing any information to evaluate and determine the equality of any proposed product.  Therefore, ensure that sufficient information and descriptive material is provided such as cuts, illustrations, drawings, or other information to establish clearly and precisely what product the supplier is proposing and demonstrate that the product offered meets the requirements of this specification.

Packaging PREPARATION

Do not perform packaging and shipping of GBs until shop acceptance testing and inspection have been performed and the results approved by LANL.  Prepare and package GBs and associated components to prevent damage during shipping and handling.  Use particular care to ensure that the surface finishes, cleanliness, dimensional stability, and overall integrity of the GBs achieved during fabrication are not affected during shipment.

Install windows loosely on the GB or ship separately.  Electrical connector installation is optional, depending upon vulnerability.  Seal GB openings with temporary covers or other protection to exclude dirt and prevent damage to openings of fittings or sealing surfaces of gasketed openings.  Protect mating surfaces with clean plywood or cardboard covers.  If tape is used, use tape consisting of low chloride (250 ppm) content.

If GB is fabricated at an elevation different from the elevation indicated in the site conditions section of this specification, provide for a means of pressure relief in the GB during shipping.  An open service penetration on the GB with temporary dust filter will provide adequate means of pressure relief.  Do not seal GB during shipping.

PACKAGING

Packaging Procedure: Submit a packaging procedure that meets Level B requirements as described in Subpart 2.2 of ASME NQA-1.  This procedure describes the methods, material control, and inspections to be used by the supplier to perform GB packaging for shipment.  The procedure addresses the covering of GB openings, pallet and crate construction, protection of the GB sealing surfaces, and marking of the crate.  Submit this packaging procedure for approval prior to performing this work.

Pack GBs individually in totally enclosed wooden crates with pallets or provisions for handling by a forklift.  Ship GBs as complete assembled units except for support stand legs.  Fully welded support stands to be shipped separately from the GB shell.  If shipping limitations restrict complete assembly shipments, propose a recommended alternative for approval by LANL.  Separately pack components or items that may work loose or be lost in transit.

Bag or crate any separate components that are part of the order and mark accordingly to describe or identify the GB with which the components are associated.  Furnish packing material, weather protection, dunnage, and crating.

Provide desiccant bags inside GBs to prevent condensation build-up during shipping.

CRATING

Provide lumber seasoned, reasonably sound, and free from cross grain and knots that would interfere with nailing or stapling, or knots that are greater than 1/3 the width of the lumber.

Construct crates with outer framework consisting of upright and horizontal members and with additional diagonal upright and horizontal members where necessary to provide proper strength and rigidity.

Construct crates with three-way lock corners, where members will be joined with nails or staples driven into side grain of joining members.  See the following examples of three-way lock corners.

Use double nailing, stapling, or screws to fasten joining crate members.

Design and construct crates with transverse cross-members at the base sufficient in strength to protect the underside from damage by mechanical handling equipment.

EXAMPLES OF THREE-WAY LOCK CORNERS

Crate Marking: Properly and clearly mark crates on the top and four sides using a stencil.  As a minimum, provide information including LANL Subcontract number, and the actual weight of the crate and its contents.  Identify each crate or package as a part of the total order, for example “Crate #1 of 5.”

SHIPPING

Shipping shall be in accordance with Level B requirements of ASME NQA-1, Subpart 2.2, Paragraph 400 Shipping.  Provide LANL with a copy of the bill of lading concurrent with the shipment.  Properly and clearly describe the shipment on the bills of lading.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Provide final inspection and acceptance upon receipt of the shipment at Los Alamos.  Inspect the shipment as necessary to ensure that received items have not been damaged during shipment and that required items and supporting documentation have been received.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

RECEIVING

Final Inspection and Acceptance: LANL will inspect the shipment as necessary to ensure that received items have not been damaged during shipment and that required items and supporting documentation have been received.

Warranty

Guarantee the GBs and GB equipment at design conditions and warrant that materials and workmanship or apparatus supplied, are in accordance with Subcontract document requirements.

site conditions

GBs will be installed at an altitude of 7,500 feet above sea level at Los Alamos National Laboratory, located in Los Alamos County, NM.

- products

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

An approved vendors list shall be established for GB vendors in accordance with the appropriate Quality Assurance Program Plan.  Authors shall select vendors from the approved vendors list.  This approved list is the Institutional Evaluated Suppliers List (IESL).

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

- suppliers

Companies specializing in manufacturing products with suitable documented experience of performing similar work must be listed on the Institutional Evaluated Suppliers List (IESL)

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine and specify the type of material to be used for the GB shell, GB support stand and other appurtenances.  Specify material based upon application and compatibility with the process and corrosive environments.  (e.g., carbon steel may be used for fabrication of the GB support stand in lieu of stainless steel, however, consider decontamination requirements outside of the GB.  Coat carbon steel support stands with a decontaminable coating.)

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

- Materials

    - Provide new materials complying with this specification section and relevant standards. Or approved equal” is always implied after a brand name, patented process or catalog number, unless “no substitution allowed” is specified.  The Subcontractor may substitute any brand or process approved as an equal by the specifying Architect/Engineer with approval.

For GB shell materials use stainless steel type [304, 304L, 316 or 316L]. Provide CMTR’s with all SS used in the GB fabrication.

Stainless Steel Sheet

The phrase “300-series stainless steel” is used throughout this specification section.  300-series stainless steel refers to types 304, 304L, 316, or 316L stainless steel.  Type 304L may be substituted for type 304, and type 316L may be substituted for type 316.  Do not substitute type 304L or type 316L material when it is specified on the Drawings.  If 300-series stainless steel is specified, one of the four types of stainless steel as previously stated may be used at the discretion of the supplier.

Fabricate GBs from 300-series stainless steel, as specified on the Drawings.  Use 7-gauge sheet stock with a No. 2B mill finish on both sides conforming to ASTM A240 and A480 for shell material.

Fabricate 16-gauge sheet used to cover the lead sheet on shielded GBs from 300-series stainless steel as specified on the Drawings.  Provide sheet with a No. 2B mill finish on both sides.

It is recommended that the supplier purchase stainless steel sheet with a protective coating on both sides to protect the mill finish.

Stainless Steel Plate

Do not use plate stock except where so specified in the Drawings.  When plate stock is required for use as a reinforced floor or as mounting pads for equipment, use the same grade of stainless steel as the shell material per ASTM A240.  Polish the plate stock to the surface finish required by § 3.4. Note: Thickness of plate shall be limited to 10% thinning due to grinding/polishing.

Stainless Steel Bars and Shapes

Provide structural bars and shapes used in GB fabrication in the same grade of stainless steel as the GB shell material, meeting the requirements of ASTM A276.

Where approved by LANL, equivalent shapes may be fabricated by bending the appropriate sheet or plate stock.

Stainless Steel Forgings

Provide forged couplings and other forged pieces used in GB fabrication as the same grade of stainless steel as the GB shell material.  Ensure forgings meet the requirements of ASTM A182 or ASTM A479.  Castings are not permitted unless specified in the Drawings.  Castings are permitted for 150# rated pipe elbows.  Couplings rated for 3000# may be used in lieu of couplings rated for 150# when the couplings are unavailable in the same material as the GB shell.

Stainless Steel Structural Shapes

Provide stainless steel structural shapes including channel, I-beam, and L-angle per the chemical and physical requirements of ASTM A276.

Stainless Steel Structural Tubing

Ensure square and rectangular tubing used for the support stand meet the chemical and physical requirements of ASTM A554.

Carbon Steel

Ensure carbon steel for the GB support stand conforms to the chemical and physical requirements of ASTM A36, ASTM A572, or A588 standards as specified on the Drawings. Provide CMTRs for the carbon steel used for support stands components.

Welding Materials

Use filler materials that conform to ASME B&PVC Section II, Part C and as follows:

| Base Material | Filler Material |
| --- | --- |
| Type 304 SS | ER 308 |
| Type 304L SS | ER 308L, or ER 347 |
| Type 316 SS | ER 316 |
| Type 316L SS | ER 316L |
| ASTM A-36 or equivalent | E70xx, ER70S-xx, or E70T-xx |

Provide CMTRs or CoCs for welding filler materials used in the fabrication process as part of the QA Document Package prior to start of welding activities.

Devcon Stainless Steel Putty, part no. 10271 may be used for exterior seal welds. However, the Devcon cannot be used for structural integrity of the part.  Use is restricted to eliminating contamination traps between adjacent parts on the exterior of the GB only.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine and specify the appropriate lead thickness in accordance with radiation shielding analyses.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Lead

Ensure lead used in the fabrication of applicable portions of shielded GBs meets the chemical requirements of ASTM B749.  Submit a CMTR or CoC as part of the QA Document Package.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine and specify the type of gloveport to use for GBs.  Determine whether gloveports will be bolted into the GB shell or windows.  Determine whether gloveports will be welded into the GB shell.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Gloveports

Use push-through gloveports from Central Research Laboratories (CRL) of the round or oval shape for GBs.  Provide weld-in or clamp-in gloveports per the Drawings.

Provide gloveports of the [rolled and welded type or CRL push-through type] for hoods, or other approved designs.

Provide o-ring materials for bolt-in, push-through gloveports of type [neoprene, buna-N, Viton, silicone].

GB Gloves

Refer to Section 11 5311.14 – Glovebox Gloves.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine and specify type of viewing window material to use for GBs.  GB windows should be constructed of wire glass, fire-rated glass or laminated safety glass.

If laminated safety glass is used in locations where radiation levels may be high enough to cause yellowing of the plastic, a cerium additive should be specified to prevent yellowing.

If the GB atmosphere or operations require an alternate window be used, fire-retardant treated polycarbonate may be used; fire-retardant treated polycarbonate may be used sandwiched with noncombustible material.

Note: Lexan® is considered a combustible product and may have an impact on facility combustible loading calculations.  Use of Lexan and impact on facility combustible loading shall be considered prior to specification and use for GB windows.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Windows

Provide viewing pane material of [safety glass, leaded safety glass, or polycarbonate resin].  Provide glass thickness of [1/4-inch, 3/8-inch, 1/2-inch, 5/8-inch].

Provide window materials with CMTR.  CoC may be provided in lieu of CMTRs with prior approval from LANL.

Safety Glass: Provide safety glass with two equal layers of annealed plate joined by a 0.030- to 0.050-inch laminate layer of polyvinylbutyral (PVB).  Provide safety glass with a density of 2.5 g/cc and a refractive index of 1.52.  Provide glass meeting the requirements of ASTM C1036 and ASTM C1172.

Leaded Glass: Use lead impregnated safety glass where approved by LANL for shielding window applications.  Coat lead glass with float glass on both sides for protection from scratching and chipping.  Lead glass density should not exceed 5.2 g/cc to avoid yellowing.

Polycarbonate Resin: Provide fire retardant grade material with silicate coating.

Chemically-Strengthened Glass: Provide chemically strengthened, laminated safety plate glass, Toroglass® from Hot Cell Services.  Provide plate glass with a density of 2.5 g/cc with an index of refraction of 1.52.

Provide window gaskets per Section Q.

Fasteners

Provide fasteners with CMTR.  CoC may be provided in lieu of CMTRs, when CMTRs are not available from the mill or subtier supplier.  Where approved by LANL prior to submittal, CoCs may be submitted in lieu of CMTRs.

Use UNC-series threads with screws, fasteners and components in accordance with ASME B18.2.1with mating threads.

Provide bolts and cap screws with grade marks.

Implement provisions to ensure that materials used or supplied are not counterfeit or of other suspect origin.  Pay particular attention to high strength bolting materials (grade 5 and 8 strength) and pipe fittings.  Reference Attachment 3 for further guidance on suspect fasteners.

Stainless Steel

Unless otherwise specified on the Drawings, use stainless steel for nuts and washers used in the GB fabrication.

Use stainless steel fasteners type 304, 304L, 316 or 316L.

Stainless Steel Bolts and Hex Head Cap Screws: ASTM F593.

Stainless Steel Nuts: ASTM F594.

High Crown Acorn Nuts: AISI 300-series

In the event that fasteners are not readily available meeting the requirements of ASTM F593 and ASTM F594, then fasteners meeting the requirements of ASTM A193 and ASTM A194, respectively, may be substituted with prior approval by LANL.

Ensure all fasteners conform to the following standards:  ANSI B18.2.1, SAE J429, and ASTM A354.

Carbon Steel

Carbon steel bolts per ASTM A307 or A325 and nuts per ASTM A563A.  Ensure all fasteners conform to the following standards:  ANSI B18.2.1, SAE J429, and ASTM A354.

Stainless Steel Weld Studs

Ensure capacitive discharge (CD) weld studs welded to the exterior of the GB are 300 series stainless steel and meet the chemical requirements of ASTM A276.  Ensure weld studs welded to the interior of the GB are of the same type stainless steel as the GB shell.  Provide CoCs or CMTRs with CD weld studs.

Electrical Feedthroughs

Refer to Section 11 5311.16 – Glovebox Feedthroughs, Hermetically-Sealed.

Gaskets and Seals

- The general ASTM D2000 format for the gasket callout is: ASTM D2000-12 M1XX*##. The resulting callout string will be ASTM D2000-12-M1BC*07” for neoprene and ASTM D2000-12-M1HK*07 for Viton. Note, the tensile strength callout of “07” is a minimum value. Table 1 below shows each of the callouts in the string and the associated tests per ASTM.

Table 1: ASTM D 2000-12 M2BC*07/ASTM D 2000-12 M2HK*07

|  | Meaning | Test |
| --- | --- | --- |
| M | SI Units | ASTM D 2000-12 |
| 1 | Grade | ASTM D 2000-12 |
| BC | Chloroprene polymer (Neoprene) | ASTM D 2000-12 |
| HK | Fluorinated elastomers (Viton, Fluorel, etc.) | ASTM D 2000-12 |
| * | Durometer; see drawing | ASTM D 2240, Shore A |
| 07 | Tensile 7 MPa, 1015 psi | ASTM D 412-06a |

  - Associated tests to be per latest version of ASTM D2000.

- Product to be shipped with the manufacturer's QA data package containing documentation discussed in the following steps. The documents must be legible and traceable to LANL's purchase order/procurement document, and reflect information in compliance with the procurement document specification. For each lot, the QA data package must include:

Manufacturer's Certificate of Conformance (C of C) stating that the supplied items conform to the LANL specifications contained in the LANL procurement document. The C of C must be signed or otherwise authenticated and dated by an authorized representative as defined in the manufacturer's quality assurance program. The certification must identify the manufacturer, be traceable to the procurement document and include the following:

- Item manufacture/cure dates

- Model numbers

- Any approved changes, waivers, or deviations from the procurement specification, as provided to them in the procurement document.

Material Transfer Devices

Provide material transfer devices including [bagports, double-door transfer systems, airlocks, input sphincter, and introductory tubes].  Attachment 1 of this section provides design details of some of these material transfer systems.

Provide double-door transfer systems from Central Research Laboratories or approved vendors.  Provide double-door transfer systems of diameter [105 mm, 190 mm, 270 mm, or 350 mm].

Provide [rolled and welded bagport, or CRL push-through bagport] for use with [30-gallon drums, 55-gallon drums].  Provide rolled and welded bagports in accordance with drawings in Attachment 1.

Provide airlock of size [XX-inches diameter by XX-inches long].  Provide airlock doors of a [counterbalance style, guillotine style, cam style or other approved design].  Provide a slide tray inside the airlock.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

For non-UL rated components, verify the acceptability of the components prior to their use by the supplier.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

UL-Rated Components

Where certified products from Underwriters Laboratories Inc. (UL) or another nationally recognized testing laboratory are available, provide them in lieu of non-certified units.  Request approval from LANL for the use of non-UL certified products.

Paint

Primer – Primer is no longer necessary, however, if required use as specified:

Provide coating materials of the type and color specified on the Drawings.  If unspecified, use a heavy-duty primer that meets the following requirements:

          - Minimum dry film thickness of 2.5 mils.

Heat resistance of 200( F.

Minimum volume solids content of 43%.

Finish

Ensure coating materials are of a type and color specified on the Drawings.  If unspecified, use a self-priming cross-linked epoxy-phenolic cured finish with a polyamine curing agent (PLASITE® 7122 VOC by Carboline®, or equal).  For self-priming finishes they must meet the combined requirements of both Primer and Finish. Otherwise, provide a finish that meets the following requirements:

          - Pearl gray color.

Minimum dry film thickness of 5 mils.

Heat resistance of 200( F.

Anodize all fabricated aluminum per MIL-A-8625, Type II, Class 1 (non-dyed).

- As-Built Drawings:

    - Prepare as-built drawings to reflect the GB, as fabricated.  Provide as-built drawings consisting of either CADD-generated drawings or a redline markup of the fabrication Drawings.  Submit as-built drawings as part of the QA Document Package.  Document the following in as-built drawings:

Modifications and Deviations: Reflect modifications and deviations to the Drawings that have been approved by LANL and subsequently implemented by the supplier.  Do not note dimensional exceptions fabricated within the tolerances of the Drawings.

Weld Locations on Weld Map: Indicate GB shell (primary confinement) weld locations on the weld map drawings.  Do not depict other welds for stiffeners, stands, attachments, and outer skin.  Provide numbering of welds and identification of welders.

Material Identification: Identify material and heat number of each sheet (or piece – plate) of material used in the GB shell weldment on the weld maps.

execution

- general fabrication requirements

Perform cutting with mechanical shop tools, plasma arc, laser, or water jet.  Do not use carbon arc or iron powder cutting on stainless steel.

Ensure cut or raw edges are deburred and smooth to the touch.

Ensure shell bends have an inside radius of 5/8 inches with other bends having a minimum inside radius equal to the thickness of material, unless otherwise specified on the Drawings.

Chase coupling threads with tap after the coupling is welded into place.

Use wire brushes made of stainless steel.  Ensure grinding wheels and wire brushes are new or previously used only on stainless steel.

Ensure temporary carbon steel clamps, supports, braces, and fixtures used during fabrication are not welded directly to, or come into direct contact with, any stainless steel surfaces.  Do not use galvanized steel clamps or fixtures.

Clean press brake dies with solvent before use in forming stainless steel materials.  Ensure carbon steel parts of the press brake that will come in contact with the stainless steel material are masked or covered to control carbon contamination.

Provide initial weld fit-up inspection after the GB is formed and tack welded in place but before final welding.  This inspection will also include the preliminary critical dimensional inspection, weld joint verification, weld map verification and GB orientation.  Note: Critical dimensions shall be provided by Design Engineer for verification. Provide LANL with a seven (7) working day advance written notice so that a LANL representative may witness the activity.

- Dimensional Control

    - Ensure critical dimensions and acceptance tolerances specified on the Drawings apply to the finished GB or component.  Indicate compliance with all dimensions by developing and submitting a Dimensional Control Map.  Verify all dimensions on drawings and document measured dimensions on Dimensional Control Map.  Also document actual surface flatness measurement around sealed openings.  Submit Dimensional Control Map and Weld Map verification with QA Document Package.

    - Ensure flatness of the GB cutout surfaces by:

Providing a dimensional inspection report by taking measurements at least every 6 inches around the perimeter of the cutouts with respect to the entire plane. The measurements shall be documented on a flatness grid map numerically or graphically. This inspection report (flatness grid map) shall be submitted with the QA Document Package.

Dimensional inspection for window flatness includes planar inspection, to consist of measuring across the opening for flatness of adjacent sides and opposite sides of the window cut out.

Ensure tooling or equipment (any type of coordinate measuring machine) used for determining surface flatness is calibrated and capable of measuring to 10 percent of the minimum flatness tolerance defined in the Drawings.  Flatness of cutout surfaces shall be measured in a linear side as well as the complete plane of the opening.

Submit to LANL for their review and approval the methodology/procedure to be used to verify cutout surface flatness prior to fabrication.  The dimension mapping procedure should have the actual dimensions shown on the drawing, including the out of tolerance dimensions circled in red.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Delete the “Lead Placement” section for non-shielded GB specifications.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

- Lead Placement

    - Do not glue lead to the GB, for shielded GBs requiring lead sheet application.

Ensure there are no gaps in the lead joints.  Eliminate gaps by filling with lead wool, fusion welding, or by peening.

Ensure lead applied in multiple layers have staggered joints.

Ensure gaps between lead sheets and structural framing are no larger than 1/8 in.

Pack lead wool into gaps between lead sheets and structural framing.  Refer to LANL Drawing 26Y-202001.

Provide LANL with a seven (7) working day advance written notice so that a LANL representative may inspect the lead fit-up, prior to being covered with cladding.

- Surface Finishes

    - In order to preserve the original finish of the stainless steel sheet material, exercise care to prevent scratching, abrading, nicking, and denting during receiving, storage, fabrication, and handling.  Preserve the original protective coating as long as possible.

After fabrication is completed and before testing and inspection, clean, de-scale, and degrease GBs and associated components.  Do not paint stainless steel surfaces, interior or exterior, unless specified on the Drawings.

Surface finishes for GBs are specified below and on the Drawings.

Exterior, interior and cladding: Provide exterior surfaces with a 2B mill finish except welds or damaged surfaces.

Welds: Grind and polish welds parallel to the weld to a 32-microinch-roughness height (arithmetical average) finish and blended to the adjacent material.  Limit grinding and polishing to the zone disturbed by the welding not exceeding a 2-in. width.  Ensure welds covered by shielding are ground so that there is no unevenness in the shielding.  Perform liquid penetrant testing on welds covered by shielding.  Weld sizes and base material must be maintained after grinding and polishing.

Damaged Surfaces: Polish damaged surfaces to a 32-microinch-roughness height (arithmetical average) finish.  Limit the extent of refinishing to the immediate damaged area.

Plate: When plate stock is required, polish it to a 32-microinch-roughness height (arithmetical average) finish.  Grind plate surfaces to remove pickled finish.  Liquid penetrant plate surfaces to locate pits.  Repair pits with weld.  Grind repair welds and liquid penetrant test again, until pits are removed.  Grind and polish repairs to blend with the surrounding material.

Appurtenances: Finish appurtenances, such as, but not restricted to doors and door hardware, shelves, brackets, and machined components to a 32-microinch-roughness height (arithmetical average).

Openings: Polish a band 1-inch wide around window and panel openings to a 32-microinch finish such that the grain is parallel to the edge of the opening.

Provide ALL fabricated products free from burrs or sharp edges.

- Welding

    - Welding Procedure Specification (WPS): Use welding procedures specifying standard stainless steel (P8 to P8) WPS per ASME B&PVC Section IX.  Use WPSs that cover the entire range of material thicknesses being welded.  Provide welding procedures that address weld repairs.  Submit welding procedures (WPS) with supporting PQRs to LANL, for approval prior to fabrication.

Welding Procedure Qualification Record (PQR Procedure Qualification Records (PQRs) used to support required WPSs shall meet the requirements of ASME B&PVC Section IX.    Submit PQRs with the appropriate WPS to the LANL for approval, prior to fabrication.

Welder Performance Qualification Records (WPQR): Ensure welders, welding operators, and tack welders are qualified in accordance with ASME Boiler and Pressure Vessel Code (B&PVC) Section IX.  Use welders who have successfully performed welder qualification tests for all positions.  Use welders for welding pipe and tubing who have successfully performed welder qualification tests in the 2G&5G or 6G position. Welder continuity records shall be submitted with WPQRs performed more than 180 days prior to fabrication. Submit records to LANL prior to fabrication.

Welding Processes:

Note:  Unless otherwise stated on the Drawings, do not weld carbon steel to stainless steel.

GTAW is the preferred welding process for 7 gage or less stainless steel.  Other welding processes or combination of processes may be used if specifically approved by LANL. Submit welding procedure specifications (WPSs) and supporting PQRs to LANL for approval prior to fabrication.

Use certified shielding gas as specified in the supplier’s welding procedure specifications (WPS). Provide Certificate of Analysis (C of A) per AWS A5.32.

Cleaning before Welding: Prior to welding, interior and exterior surfaces, remove dirt, scale, corrosion, dust, grease, oil, water, oxidation, or foreign material.  Do not use a carbon steel brush for cleaning. Use only new or not previously used (on carbon steel) stainless steel brushes on stainless steel fabrications.

Weld Joint Design:

Material 1/8 in. or thinner may be fusion welded without joint spacing between work pieces. Bevel material of 1/8 in. or thinner if required to obtain the specified full penetration weld.

Bevel material with thickness greater than 1/8 in., but less than 1/2in. to provide a 1/16 in landing with either a 60( V-type or 75( U-type butt joint.

On tee and corner joints or material 1/8-in thick or greater, provide one work piece with a 45( bevel with a 1/16-in. land.

Welds for Butt Joints

Provide welds for butt joints with full penetration and with a uniform transition from the joined materials into the weld deposit.  Ensure the welds are free of undercutting and un-fused overlap of the weld deposit.

Ensure the width of the finish weld layer of butt welds are held to a minimum and do not exceed the width of the weld groove by more than 1/16 inch on either side.

Fillet Welds

Provide fillet weld surfaces with a uniform transition from the weld to the base material.  Provide welds free of undercutting (≤ 0.010”) and un-fused overlap of the weld deposit.

Ensure the minimum permissible length of each leg of a fillet weld is equal to the required size of the weld as called out on the procurement drawing, or equal to the thickness of the lighter section being joined if no size is specified.  Weld sizes shall be maintained after grinding and polishing. Seal welds are sufficient for sealing with no minimum size requirements. Seal welds shall not be used to hold members in place.

Weld Defects

Provide welds free from defects including the following defects and conditions:

Cracks of any description in the weld or base metal.

Crater cracks.

Slag inclusions, tungsten inclusions, oxide inclusions, or porosity.

Cold lap in the deposited weld metal.

Overlap of weld metal on the base metal.

Undercutting at the edge of the welds ≥ 0.010”.  Ensure no part of the finished face of weld in the area of fusion of welded joints lies below the surface of the base metal adjoining the weld.

##### Insufficient fill-in butt welds (CJP or PJP) below the surface.

Unfilled weld craters or shrinkage cavities.

Evidence of damage to the weld metal through oxidation.  Oxidation is defined as scaling of the metal that cannot be removed or restored to a bright metal by wire brushing with a stainless steel brush.  Heat discoloration is not considered oxidation.

Weld spatter.

Arc burns or scars on the base metal caused by striking or dragging the welded arc across the base metal.

##### CJP butt welds with less than 100% penetration.  CJP butt joints normally welded from one side only may be welded from both sides to obtain 100% penetration, provided that no excessive distortion occurs in the sections being joined and the joint design is included in the approved WPS

Intermittent or skip welding, except on exterior stiffeners and unless specified as such on the Drawings.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Note that welding of GB stiffeners to GBs can cause warping of the shell when heat is applied near other welds and openings on the GB.  Location of stiffeners and sizing of welds for stiffeners needs to be considered during design of the GB.  Suggest the use of stitch welding to attach stiffeners to GB shells.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Stud Welding

Use only capacitive discharge (CD) weld studs, unless a stud size 3/8-inch in diameter or larger is specified on the Drawings.

##### Stud welding shall be performed in accordance with an approved WPS. The WPS shall be developed in accordance with AWS D1.6 and the following requirements.  Submit the welding procedures WPS and appropriate PQR to, for approval prior to fabrication.

##### Welding Procedure Qualification Record (PQR): Procedure Qualification Records to support required WPSs shall meet the requirements of AWS D1.6 and the following requirements. Each diameter and or change of geometry of the stud base is cause for re-qualification.

##### PQR shall consist of ten studs welded consecutively using the recommended procedure and settings in each position to be welded (1S, 2S & 4S), 30 studs total, with the same surface geometry to be welded.

##### All 10 studs from each position shall be visually examined for full 360o flash.

##### Bend test: 5 studs for each position shall be tested by being bent 900 from original axis. Test is acceptable if stud is bent 900 and no fracture occurs in the weld.  Fractures occurring in the parent metal or the shank of the stud shall be acceptable.

##### Torque or Tension test: The 5 remaining studs for each position shall be torque tested if threaded, or tension tested if unthreaded. Torque testing shall use the torque testing apparatus that meets the requirements of Figure #1. The stud application shall be considered qualified if all test specimens are torque tested to destruction without failure in the weld. Tension testing shall use the tension testing apparatus that meets the requirements of Figure #2. The stud application shall be considered qualified if all test specimens do not fail in the weld.

##### A change in power supply model or stud gun model shall require re-qualification.

##### Female studs to be tested are determined by the nominal thread size and not the outside diameter or base size.

##### Studs shall not be fillet welded.

##### Production welding – At least daily

##### Provide base of male and female studs flush with the sheet to which they are welded.  Ensure the axis of studs is perpendicular to the plate to which they are welded within 2(.

##### Stud welding verification shall be performed at the beginning and end of each shift or whenever there is a change in series of male or female studs. Stud welding verification before and after fabrication: Prior to production welding, four studs shall be welded using the same settings, position, technique and thickness (± 25%) as those to be welded in production. All studs are to be visually examined to ensure a 360o flash. Two studs shall be bent to 300 from original axis. Acceptance shall exhibit no cracks in the weld. The remaining studs shall be torque tested using a torque testing apparatus that meets the requirements of Figure #1. Acceptance criteria shall be in accordance with Table #1. If studs are not threaded they shall be bend-tested also. If a failure occurs in testing, the machine may be adjusted and the test repeated.  After fabrication the same bend test shall be performed with two studs, on similar material, at the end of production to assure process consistency. If a failure occurs in this testing, the production studs shall be torque tested in accordance with Table #1.

##### Note: Female studs cannot be bent, they must be torque tested.

##### Remove all weld spatter.

##### Submit test reports for each series of stud welding to LANL.  Perform the setup and testing as described at the beginning and end of each shift or if there is a change in diameter of stud or WPS being used.

##### Torque testing of female studs is determined by the nominal thread size and not the outside diameter or base size.

###### Figure #1

###### Table #1

###### Figure #2

### Production Stud Test Reports

#### Provide test reports for each series of stud welding operations.  Ensure these tests comply with the supplier’s stud test procedure.  Include the following information (as a minimum) in the test reports:

##### GB identification.

##### Date of test.

##### WPS used.

##### Name and signature of stud weld test operator.

##### Stud welding power supply and stud gun make, model and settings.

##### Torque values at failure or tensile test values at failure.

##### Failure mode of each of four test studs.

## Welding GB Stands

    - Welding Procedure Specification (WPS): Use WPSs meeting the requirements of AWS D1.1, D1.6, or B2.1 Standard for Welding Procedure and Performance Requirements, as applicable. Provide Welding Procedures Specifications that also address weld repairs. Submit the welding procedures (WPSs) with supporting PQRs to LANL, for approval prior to fabrication.

    - Welding Procedure Qualification Records (PQRs) used to support required WPSs shall meet the requirements of AWS B2.1 Standard for Welding Procedure and Performance Requirements or AWS D1.6 for stainless steel welds, and/or AWS D1.1 for carbon steel welds.

    - Welder Performance Qualification Records (WPQR): Ensure welders, welding operators, and tack welders are qualified in accordance with AWS B2.1 Standard for Welding Procedure and Performance Requirements, or AWS D1.6 for stainless steel welds, or AWS D1.1 for carbon steel welds. Use welders who have successfully performed welder certification tests in the 3G and/or 4G for plate, or 2G & 5G, or 6G position for pipe/tubing. Provide WPQRs for personnel performing welding operations on the GBs. Welder continuity records/logs shall be submitted with WPQRs qualified more than 180 days prior to fabrication. Submit records to LANL for approval, prior to fabrication.

    - Welding Processes:

      - Unless otherwise stated on the Drawings, do not weld carbon steel to stainless steel.

      - GB stands may be carbon steel or stainless steel.  GTAW is the preferred welding process for stainless steel. Other welding processes or combination of processes may be used if approved by LANL. Use (CMTR) shielding gas as specified in the supplier’s welding procedures.

    - Cleaning Before Welding: Prior to welding, remove dirt, scale, corrosion, dust, grease, oil, water, oxidation, or foreign material from surfaces to be welded.  Do not use carbon steel brushes for cleaning stainless steel. Use only new or not previously used (on carbon steel) stainless steel brushes on stainless steel fabrications.

    - Weld Joint Design:

      - Material 1/8 in. or less may be fusion welded without joint spacing between work pieces.  Bevel material of 1/8 in. or less if required to obtain the specified full penetration (CJP) weld.

      - Bevel material with thickness greater than 1/8 in., but less than 1/2 in., to provide a 1/16 in. root face with either a 60( V-type or 75( U-type butt joint.

    - Welds for Butt Joints:

      - Provide welds in butt joints with full penetration (CJP) with a uniform transition from the joined materials into the weld deposit.  Ensure the welds are free of undercutting (≤ 0.010”) and un-fused overlap of the weld deposit.

      - Ensure the width of the finished weld layer of butt welds are held to a minimum and do not exceed the width of the weld groove by more than 1/16 inch.

    - Fillet Welds:

      - Provide fillet weld surfaces with a uniform transition from the joined material into the weld deposit.  Provide welds free of undercutting (≤ 0.010”) and un-fused overlap of the weld deposit.

      - Provide fillet welds, unless otherwise specified, symmetric with respect to the components they join.

      - Ensure the minimum permissible length of each leg of a fillet weld is equal to the required weld size as specified on the Drawings, or equal to the thickness of the thinner section being joined, if no size is specified.  Seal welds have no minimum size requirements unless specified otherwise by the engineer.

    - Weld Defects:

      - Provide welds free from the following defects exclusive of material conditions:

        - Cracks of any description in the weld base metal.

        - Crater cracks.

        - Slag or tungsten inclusions or porosity.

        - Cold lap in the deposited weld metal.

        - Overlap of weld metal on the base metal.

        - Undercutting at the edge of the welds in excess of 0.010”.  Ensure no part of the finished weld face lies below the surface of the base metal adjoining the weld.

        - Insufficient fill in butt welds below the work surface on either side of the welded seam.

        - Unfilled weld craters or shrinkage cavities.

        - Evidence of damage to the weld metal through oxidation.  Oxidation is defined as scaling of the metal that cannot be removed or restored to a bright metal by wire brushing with a stainless steel brush.  Heat discoloration is not considered oxidation.

        - Weld spatter.

        - Arc strikes, on or adjacent to the weld, or on the base metal caused by striking or dragging the electrode across the base metal.

        - CJP butt welds with less than 100% penetration.  CJP butt joints normally welded from one side only may be welded from both sides to obtain 100% penetration, provided that no excessive distortion occurs in the sections being joined and the joint design is included in approved WPS.

- stainless steel cutting

    - Air arc cutting of stainless steel is not permitted.

If plasma cutting is used, grind cut surfaces to remove oxide surfaces. Use only aluminum oxide grinding wheels, new or not previously used on carbon steel.

If stainless steel is nibbled, grind points to smooth base metal.

Shearing of material is acceptable.

Laser or water jet cutting is preferred, if available.

- Chloride content control

    - Exercise control during stages of fabrication to minimize exposure of stainless steel to contaminants, in particular any chloride that might cause stress-corrosion cracking.

Avoid chloride-bearing compounds; however, if used, remove them completely by thorough cleaning with chloride free, demineralized water, lint-free rags.  Use compounds, liquids, or markers that come into contact with stainless steel surfaces with no more than 250 ppm by weight chloride.

Submit Material Safety Data Sheets or independent lab test reports as part of the QA Document Package for cleaning solvents, tape adhesive, and marking pens.

- PAINTING

    - Scope: Paint carbon steel in accordance with this specification and as required by the Drawings.  Do not paint stainless steel components.

Surface Preparation

Clean and properly prepare surfaces to be coated before any coating is applied.  Prevent rusting and/or contamination of cleaned or primed surfaces.  Coat the cleaned surfaces the same day cleaning is done and before detrimental corrosion or recontamination occurs.  Remove oil, grease, and other contaminants by solvent cleaning in accordance with SSPC SP-1 before any mechanical cleaning.

Use abrasives for blast cleaning that are clean and dry.  Select abrasives to provide a proper surface profile for the subsequent priming materials.  Ensure air pressure supply lines for blasting have effective and proper moisture and oil trap filter devices.  Perform blast cleaning in accordance with SSPC SP-6.

Application

Apply paint to dry, clean, adequately prepared surfaces, in accordance with manufacturer’s instructions.  Properly cure each coat of paint according to manufacturer’s instructions before applying additional coats.

Ensure paint containers remain closed until required for use.  Mix paint before use in accordance with manufacturer’s instructions.  Provide agitation during application where specified by the manufacturer.

If the total dry film thickness is not obtained in one coat, apply additional coats until the specified thickness is provided.

Apply coats in such a manner as to produce a film of uniform smoothness.  Pay special attention to crevices, weld lines, corners, and edges to obtain the required thickness.

Follow the manufacturer’s instructions for thinning, mixing, handling, and applying the products as part of these specifications.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Determine and specify appropriate information for filling in GB labels.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

- lABELING

    - Attach an identification nameplate to each GB with the GB number and procurement drawing number on the first line and the grade of stainless steel and the weight of the GB on a second line as shown in the example below.

| 2001-GB-100 | 55Y-999999 |
| --- | --- |
| 304L | 1200 LB |
| Date Fabricated: XX/XX/XXXX |

Provide nameplate manufactured from 16-gauge stainless steel of the same grade as the GB shell.  Provide nameplate with dimensions of 2-in. high and 5.5-in. long, fusion welded to the exterior of the GB.   Stamp lettering into the nameplate with 3/8-in. high letters and numbers. Fill stamping with black ink to make the letters legible.  Place the nameplate at the lower right corner of the front of the GB if possible.

A supplier identification nameplate may be attached to the GB above the welded nameplate.  Ensure the supplier nameplate does not exceed eight (8) square inches.

- Assembly

    - Conditions

Assemble GBs in a clean, dust-free area of the supplier’s shop.

Assemble multiple adjoining GBs together in an upright position on a level surface.

Check to ensure the entire length and height of the GBs are parallel and perpendicular in accordance with the Drawings.

Support Stand

Bolt the support stand frame to the GB to stabilize the GB structure before shop acceptance testing.

Completely fabricate, assemble, inspect, and paint the support stand before attachment to the GB.  For support stands requiring field welds, do not paint support stand.  Do not paint stainless steel support stands. Provide temporary GB support where stand configuration prevents complete shop assembly.

Keep support stand attached to the GB for subsequent fabrication operations, inspections, tests, packaging, and shipping.

The support stand or support stand legs, after being attached for verification to dimensions and tolerances specified by the Drawings, may be removed during GB shipping.

Use the support stand frame as a template for determining stud location on the GB.

- CLEANING

    - Procedure

Submit, for review and approval, a cleaning procedure describing the methods, materials, controls, and inspections to be used to perform GB-cleaning operations.

Provide procedures that address cleaning GB surfaces to remove dirt, oils, and marking pen ink.  Provide procedure that also includes a specification of the solvents and/or detergents that will be used.

Clean both interior and exterior surfaces by removing weld spatter, oil, grease, markings, from pens and dyes, shop soil, and visible rust.  Use cleaning methods that do not introduce iron or chloride contamination.  Methods may include cleaning by hot water spray or solvent wiping.  Submit the procedure to LANL for approval, prior to fabrication.

Ensure the cleanliness of the GB meets the approval of LANL at the time of the final inspection.

Detergent

If a detergent is needed to ensure thorough cleaning, use a detergent that is low in chloride.  Use fresh water for final wash and rinse.  Ensure the detergent, wash, and rinse contains less than 250-ppm chlorides.  After the water rinse, dry inside surfaces use heat, lint-free cloth, or other means to ensure cleanliness.  If heat is used for drying, ensure the final rinse water is fully softened, low chloride water with less than 250-ppm chloride.

************************************************************************************************************

When the GB will be subject to live loads such as monorails or hoists, establish additional testing and inspection requirements to verify structural/confinement integrity of the GB under all live loads and dead loads (pressure differentials during helium leak testing only test one specific dead load).

*************************************************************************************************************

- SHOP ACCEPTANCE TESTS

    - Perform the following:

Perform leak testing required by this specification.

Document torque of all fasteners on a torque map prior to performance of any shop acceptance leak testing. Indicate on the torque map the location of fasteners on the GB, torque applied to each fastener, and the calibration data for the torque wrench.  Provide the torque map as part of the QA Document Package.

Perform non-destructive examination including liquid penetrant testing and radiography where required by this specification.

Perform visual inspection of welds.  Visually verify that the welds show complete fusion of the weld metal and base metal, with no evidence of burning through, and free from surface cracking or porosity.

Perform dimensional inspection (critical dimensions only) required by this specification.

Perform surface finish inspection required by this specification.

Perform additional tests and inspection required by this specification.

Provide the test location, equipment, and instrumentation of certified accuracy and any supplementary temporary connections and auxiliary parts necessary to fully execute the tests.

Provide test personnel qualified to conduct, record, and verify test results.

Provide LANL with a seven (7) working day advance written notice of shop acceptance tests.

Submit the test results as part of the QA Document Package in accordance with Section 01330.

- nON-DESTRUCTIVE EXAMINATIONS PERSONNEL (NDE) CERTIFICATIONS

    - Provide NDE personnel performing leak testing and liquid penetrant testing operations certified in accordance with the requirements of ASNT SNT-TC-1A.

Unless witnessed by LANL, provide test reports with signatures by personnel who either performed or witnessed the test and who hold either Level II or Level III certification.

Provide NDE certifications for personnel performing or witnessing the following non-destructive testing inspections:

Helium Leak Testing; Level II minimum

Pressure Change Management Testing; Level II minimum

Liquid Penetrant Test

Radiography

Submit the NDE certifications for test personnel for approval prior to testing.

- LEAK TESTing

    - Perform leak tests on GBs, with the exception of open-front boxes. Perform preliminary/in-process leak tests and repair welding prior to the shop acceptance testing. Perform the final leak tests as part of the shop acceptance tests, witnessed by a LANL representative. Perform the helium leak tests in accordance with ASTM E498 or ASTM E499. If a Global Leak Test is required, it is to be performed following AGS-G004 (Section 6.2 or 6.3) or the applicable portions of ASME AG-1 (Articles TA-1000, TA-2000, TA-3000, TA-4000, and Mandatory Appendix TA-III).

      - The fundamental leak test requirement is a single Helium Leak Test with a maximum detectable leak of 1x10-6 std cc/sec. If the required detectable leak cannot be attained using methods described (i.e., repair and re-test), then Paragraph 3.15.A.2 below must be followed.

      - The optional leak test methodology is to perform a Global Leak Test followed by a Helium Leak test. If using Paragraph 3.15.A.1 for Acceptance testing and a detectable leak of 10-6 std cc/sec cannot be attained, then this option may be used. This optional methodology may be performed as a requirement for acceptance versus Item 3.15.A.1 above per AGS-G006 that shows a leak rate of less than 0.3% of the glovebox volume per hour when tested at an internal enclosure pressure of plus or minus 4 inches water column for a time period of not less than 15 minutes. The Helium Leak Test that follows the Global Leak Test must have a maximum detectable leak of 1x10-4 std cc/sec when tested at an internal enclosure pressure of 4 inches water column. Documentation of pressure as a function of time shall be recorded with the accompanying temperature and atmospheric pressure and be included as a Report as part of the QA Document Package per the requirements in 3.15.A.

Procedure

Submit a leak test procedure to LANL for approval, prior to testing.  Include in the procedure requirements set forth in this specification.

Describe the methods, materials, controls, and inspections used to blank off GB openings in the procedure.  The requirements for creating a helium atmosphere within the GB and how leak detection operations are accomplished are described in this procedure.  Document purge method with calculation or use of oxygen monitoring.

Test report forms are also a part of this procedure.

Assembly: Assemble the GB as much as practical.  Provide a GB assembly including windows, supplied connectors, and filter housings, but not gloves.  Seal openings in the GB shell such as gloveports and connecting rings externally with gasketed blank-off plates.  Provide blank-off plates and test coverings for openings.

Pressure: Perform testing at a test pressure of 4 in. of water gauge pressure.  To prevent pressurization above 6 in. of water gauge pressure, protect the GB during the test using a pressure relief device.

Helium Atmosphere

Provide atmosphere within the GB during testing with at least 90% helium.  Establish the required atmosphere by using either of the following methods:

Purging Method: Connect a helium gas supply to a penetration in the top of the GB.  A penetration panel coupling may be used or a fitting may be placed in a window blank-off plate.  Locate a vent near the bottom of the GB.  Flow gas from the top to the bottom of the GB.  Relate the purge flow and time by the following equation,

T = 2.3 V/Q,

where T is the time required for the purge, Q is the helium flow rate, and V is the GB volume.  Note: It is permissible to use the balloon method combined with the purge method. (V=Vbox-Vballoon).

Balloon Method: Place a weather balloon or plastic bag in the GB and fill it with helium until it expands to completely fill the GB.  Release the contents of the balloon by reaching through a glove on the GB.

Calibration: Immediately preceding the test, calibrate the leak detector against a standard that has been calibrated to a National Institute of Standards and Technology traceable leak within the last 12 months.  Use a calibrated standard leak rate between 10-6 and 10-7 standard cc/sec.

Leak Detector Probe: Move the leak detector probe along accessible joints and welds no faster than 24 in. per minute. Locate and repair any detectable leak of 10-6 standard cc/sec in a weld, bolted joint or coupling plug. Rework and repair leaks until leaks are eliminated.

If following Item 3.15.A.1 and a detectable leak of 10-6 standard cc/sec cannot be attained, then Item 3.15.A.2 may be invoked.

Shielded GBs: Test lead shielded GBs for leakage into the shielding jacket by probing the gaps around the gloveports with the leak detector.  Check neutron shielded GBs at the fill port for helium.  If helium is present, locate and repair the leaks in the shell seams.

Reports: Provide leak test reports for each GB tested other than open front boxes.  Submit final leak test reports as part of the QA Document Package in accordance with Section 01 3300.  Include the following information in the test report:

GB identification.

Date of test.

Name and signature of the test operator.

Make and model of test equipment.

Calibrated leak data.

Calibration data if detectors do not have direct leak indication.

Signature of witness.

- WELD INSPECTION

    - Visually inspect welds at the supplier’s shop during shop acceptance tests.  Perform liquid penetrant testing on primary confinement interior welds and any welds so specified on the Drawings in accordance with ASME B&PVC, Section V.

Perform liquid penetrant testing after grinding and polishing operations.  Repair and re-test detected defects.  Submit a liquid penetrant test procedure for approval prior to performing the work.  Include test report forms as a part of this procedure.  Submit the final liquid penetrant test reports as part of the QA Document Package in accordance Section 01 3300.

Radiograph 10% of full penetration Butt/Groove welds.  If defects are detected radiograph another 10% of welds and repair as necessary.  Radiography beyond 10% will be at the supplier’s expense. Acceptance shall be to Para. 3.5.1 of this spec. with porosity judged using ASME Sec. VIII Div. 1 Appendix 4.

      - A minimum of 10% of all welds shall be inspected. Welds selected for inspection shall be representative of the weldments based upon process and welder or welding operator as applicable.

      - LANL inspector or his designee will select the areas to be examination.

Progressive Sampling:

      - When sampled welds (ie welds that receive less than 10% RT examination) fail to meet the acceptance criteria given in this specification, a second sample shall be radiographed. The second sample shall be an additional 10% of representative weldments.

      - To determine the extent of the defect, the second sample shall include any weld immediately adjacent to the weld where the defect was discovered. Where practical the second sample shall include additional welds made by the same welder(s) and similar weld configuration.

      - If a defect is discovered in the second sample, all welds of the weldment shall be 100% radiographed.

Liquid Penetrant and Radiography Test Reports:  Provide test reports for liquid penetrant and Radiography inspection of welds or polished plate as required by the Drawings and this specification.  Include the following information in the test report:

GB identification.

Date of test.

Name and signature of the certified test operator.

Location and description of indications.

Description of repairs and retest.

Signature of witness.

- DIMENSIONAL INSPECTION

    - Perform dimensional inspection on the GB with a LANL witness present.  Verify that the fabricated GB is within the overall dimensions and tolerances required by this specification and as shown on the Drawings.

Use the as-built drawings to document the dimensional inspection as required by Article 2.3.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Witness the surface finish inspection on GBs at the supplier’s shop prior to shipment to the LANL site.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

- sURFACE FINISH INSPECTION

    - Perform a surface finish inspection on GBs to verify conformance of surface finishes to the requirements stated in § 3.4.  Perform inspection with a LANL representative present to witness the inspection.  Use a Surface Roughness Analyzer at random places for inspection of the weld and damaged areas.  Polish scratches and imperfections detectable by touch.

- FINAL WELD INSPECTION

  - Weld inspections to be performed by LANL Owner’s Inspectors using a weld map for review.

  - Fabricator weld inspection report review (visual, DPT, stud test etc.).

  - Airlock hydrostatic test review (report not formally requested, but test is required). Blanking ends in lieu of doors is acceptable.

  - Visual weld inspection (CWI, post grind/polish) of shell confinement welds, shell to shell welds, apron welds, and shell to apron welds.

  - Pre-coating Flatness map review.

  - Dimensional

  - Dimensional Control Map review.

  - Spot check dimensions on GB portions assemblies like windows and BSPs.

    - Check important attributes (stud size/length) of one of each configuration, more if problems are noted).

    - Inspection by assembly (for example window goes together easily, glass is centered over opening, overall sizes correct, quantity of studs is correct) is acceptable.

    - Flange of GB to GB connection fit, parallelism, etc.

  - Flatness grid map review

    - Spot check flatness of large windows, concentrate on difficult areas (upper edge, edges near significant welding, sealing surfaces crossing seams). Note that it may be OK if they are not in spec as some tuning after coating is possible.

    - Spot check flatness of doors and GB to GB flanges. Note tuning after coating is difficult.

  - Check all surfaces adjacent to areas that will be coated for burrs and sharp edges (glove cut hazards that would be hard to correct later).

  - Surface finish

    - Final inspection of backs of GB halves numerical or comparator for 32.

  - Gloveport tool fits in each gloveport.

  - Existence and placement of weld studs on GB exterior for strut support for pipe.

  - Existence and placement of weld bosses on GB interior for doors, shelves, Fire Foe.

  - All other welded features are in place (XB filter tabs, any late breaking hand/foot monitor mounts, everything that shows up on top level assy.)

  - GB labels are welded in place.

END OF SECTION

**************************************************************************

Do not delete the following reference information.

**************************************************************************

THE FOLLOWING REFERENCE IS FOR LANL USE ONLY

This project specification is based on LANL Master Specification 11 5311.10 Rev. 4, dated April 5, 2022.

Section 11 5311.10 - Attachment 1

LANL GB Detail Drawing List

| Drawing # | Drawing Title |
| --- | --- |
| 26Y-202001 | General Notes |
| 26Y-202002 | Lead Shielding & Cladding for Zippered and Bolted Gloveboxes |
| 26Y-202005 | Lead Glass Shields for Zippered and Bolted Windows |
| 26Y-202006 | Zippered Window Assembly |
| 26Y-202008 | Bolted Window Assembly |
| 26Y-202010 | Shell Penetrations |
| 26Y-202013 | Typical Gloveport Ring |
| 26Y-202014 | Bolted Service Panel Assemblies |
| 26Y-202015 | Top Access Panel Assembly |
| 26Y-202018 | Bag Ring Assemblies |
| 26Y-202019 | Removable Shelf Assembly |
| 26Y-202021 | Cooling Well Assembly |
| 26Y-202022 | Resistance Furnace Well |
| 26Y-202023 | 14” Dia. Airlock Assembly |
| 26Y-202024 | Standard Airlock Slide Tray Assembly |
| 26Y-202026 | Connector Ring Closure Cap Assembly & Connector Assembly |
| 26Y-202031 | 14” Dia. Welded Connecting Ring Assembly |
| 26Y-202032 | Sample Taking Port Assembly |
| 26Y-202034 | Reagent Transfer Device Assembly |
| 26Y-202035 | Introductory Tube Assembly |
| 26Y-202039 | 16” Square Connector Assembly |
| 26Y-202046 | Air Cylinder Mounting Assembly |
| 26Y-202047 | Air Operated Vertical Sliding Door Assembly |
| 26Y-202048 | 14” Dia. Opening Counterbalanced Door Assembly 151 R/L |
| 26Y-202049 | 16” Square Airlock Sliding Door--Hydraulic |
| 26Y-202050 | 14” Dia. Opening Air Operated Vertical Sliding Door Assembly |
| 26Y-202052 | Introductory Boxes and Hoods Exhaust Transition Piece |
| 26Y-202053 | Introductory Boxes and Hoods Upper & Lower Door Assembly |
| 26Y-202057 | 8” Dia. Exhaust (HEPA) Filter Assembly |
| 26Y-202059 | 8” Filter Housing Assembly |
| 26Y-202060 | Pressure Relief Device 161 Assembly |
| 26Y-202066 | Std. Hi-Vac System Diffusion Pump Mounting Flange Assembly |
| 26Y-202067 | Dutchman Assembly |
| 26Y-202075 | Tunnel Dropbox Transition with Firedoor Assembly (2 dr) |
| 26Y-202076 | Typical Dropbox Detail Bolt Pattern for Transition |
| 26Y-202077 | Tunnel Dropbox Transition with Firedoor Assembly 167 (1 dr) |
| 26Y-202121 | 2 x 3 Glovebox Support Stand Assembly |
| 26Y-202122 | 2 x 2 Glovebox Support Stand Assembly |
| 26Y-202123 | 1 x 3 Glovebox Support Stand Assembly |
| 26Y-202124 | 1 x 2 Glovebox Support Stand Assembly |
| 26Y-202125 | 1 x 1 Glovebox Support Stand Assembly |
| 26Y-202130 | 12” Dia. Exhaust (HEPA) Filter Assembly |
| 26Y-202131 | Neutron Shielded Glovebox Details |
| 26Y-202150 | O-Ring Gasket Seal Tubing to Valve Joint |
| 26Y-202151 | 3” and 2 1/4” I.D. Damper Assembly |
| 26Y-202152 | 4” In-Line Filter Holding Bracket |
| 26Y-202153 | Open-Front Glovebox 7 1/2” Sash Assembly |

Section 11 5311.10 - Attachment 2

List of Required Submittals

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The following list of required submittals is an example of the submittals required in association with GB fabrication.  This list is not comprehensive.  Modify the list accordingly and provide the submittal requirements for project-specific GBs with Section 01 3300, Submittals.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Note:  The (1) NQA-1 shipment copy is delivered to Facility Operations Directorate for record keeping.

| Section No: 11 5311.10 | Type of Submittal | Submittal Schedule and Number of Copies |
| --- | --- | --- |
| Section Title: Glovebox Fabrication |  |  |
| DESCRIPTION OF SUBMITTAL REQUIRED | FOR INFORMATION | FOR ENGINEERING REVIEW, COMMENT, & APPROVAL | FOR INSPECTION & ACCEPTANCE | AT PREFABRICATION CONFERENCE | PRIOR TO FRABRICATION | WITH SHIPMENT | 7 WORKING DAYS ADVANCE NOTICE | AS REQUIRED |
|  | As-Built Drawings | X |  |  |  |  | 1 |  |  |
|  | Lower Tier Services Plan | X |  |  | 1 |  | 1 |  |  |
|  | Fabrication Schedule | X |  |  | 1 |  | 1 |  |  |
|  | Bill of Lading | X |  |  |  |  | 1 |  |  |
|  | Quality Assurance Manual (Fabrication) | X |  |  | 1 |  | 1 |  |  |
|  | Fabrication and Quality Control (QC) Procedures (Include Rev # or Date): |  |  |  |  |  |  |  |  |
|  | Welding Procedure Specifications (WPS) |  | X |  |  | 1 |  |  |  |
|  | Welding Procedure Qualification Record (PQR) |  | X |  |  | 1 |  |  |  |
|  | Stud Test Procedure |  | X |  |  | 1 |  |  |  |
|  | Material Control Procedure |  | X |  |  | 1 |  |  |  |
|  | Commercial Grade Dedication Plans and Procedures |  | X |  |  | 1 |  |  |  |
|  | Shop Traveler Procedure |  | X |  |  | 1 |  |  |  |
|  | Cleaning Procedure |  | X |  |  | 1 |  |  |  |
|  | Packaging Procedure |  | X |  |  | 1 |  |  |  |
|  | Leak Test Procedure |  | X |  |  | 1 |  |  |  |
|  | Liquid Penetrant Test Procedure |  | X |  |  | 1 |  |  |  |
|  | Radiography Procedure |  | X |  |  | 1 |  |  |  |
|  | Stud Test Procedure |  | X |  |  | 1 |  |  |  |
|  | Calibration Procedure |  | X |  |  | 1 |  |  |  |
|  | Dimension mapping procedure |  | X |  |  | 1 |  |  |  |
|  | Welding & NDE Personnel List |  | X |  |  | 1 |  |  |  |
|  | Personnel Qualifications: |  |  |  |  |  |  |  |  |
|  | Welder Performance Qualification Records |  | X |  |  | 1 | 1 |  |  |
|  | NDE Personnel Certifications |  | X |  |  | 1 | 1 |  |  |
|  | Q.A. Document Package: |  | X | X |  |  | 1 |  |  |
|  | Test Reports: |  |  |  |  |  |  |  |  |
|  | Production Stud Test |  |  | X |  |  | 1 |  |  |
|  | Liquid Penetrant Test |  |  | X |  |  | 1 |  |  |
|  | Radiography Test |  |  | X |  |  | 1 |  |  |
|  | Leak Test |  |  | X |  |  | 1 |  |  |
|  | Torque Map |  |  | X |  |  | 1 |  |  |
|  | Dimensional Control Map |  |  | X |  |  | 1 |  |  |
|  | Cutout flatness grid map |  |  | X |  |  | 1 |  |  |
|  | Certified Material Test Reports: |  |  | X |  |  |  |  |  |
|  | Stainless Steel for Shell |  |  | X |  | 1 |  |  |  |
|  | Other Stainless Steel Items |  |  | X |  | 1 |  |  |  |
|  | Carbon Steel for stand |  |  | X |  | 1 |  |  |  |
|  | Lead |  |  | X |  | 1 |  |  |  |
|  | Weld Filler Materials |  |  | X |  | 1 |  |  |  |
|  | Chloride Content |  |  | X |  | 1 |  |  |  |
|  | Shop Traveler, customized for project |  | X |  | X |  |  |  |  |
|  | Certificates of Compliance |  |  | X |  |  | 1 |  |  |
|  | Supplier Deviation Disposition Request |  | X |  |  |  | 1 |  | 1 |
|  | Nonconformance Reports |  | X |  |  |  | 1 |  | 1 |
|  | Commercial Grade Dedication Records |  |  | X |  |  | 1 |  | 1 |
|  | Notification for Inspection of Weld Fit-up | X |  |  |  |  |  | 1 |  |
|  | Notification for Inspection of Lead Placement | X |  |  |  |  |  | 1 |  |
|  | Notification for Acceptance Testing & Final Inspection | X |  |  |  |  |  | 1 |  |

Section 11 5311.10 - Attachment 3

Guide to Suspect Fasteners

# QC-27       SUSPECT/COUNTERFEIT ITEMS (Jun 2011)

- A suspect item is one in which there is an indication by visual inspection, testing, or other information that it may not conform to established Government or industry-accepted specifications or national consensus standards. A counterfeit item is a suspect item that is a copy or substitute, without legal right or authority to do so, or one whose material, performance, or characteristics are knowingly misrepresented by the vendor, supplier, distributor, or manufacturer.

- SUBCONTRACTOR warrants that all items, including their subassemblies, components, and parts, tendered to CONTRACTOR shall be genuine (i.e., not counterfeit), new and unused, and conform to the requirements of this subcontract, without substitution unless otherwise provided for within this subcontract or approved in writing by CONTRACTOR prior to delivery.

- SUBCONTRACTOR further warrants that all components, parts, materials, and supplies incorporated into CONTRACTOR’S facilities or equipment by SUBCONTRACTOR, during performance of work at LANL, shall be genuine, new and unused, and original-equipment-manufacturer items, without substitution unless otherwise provided for within this subcontract or approved by CONTRACTOR in writing as suitable for the intended purpose prior to use.

- As part of the foregoing warranties, SUBCONTRACTOR also certifies that all labels and/or trademarks or logos affixed, or designed to be affixed; to items supplied or delivered to CONTRACTOR, and to certifications, affirmations, information, or documentation related to the authenticity and quality of items supplied or delivered to CONTRACTOR under this subcontract, are genuine.

- Falsification of information or documentation may constitute criminal conduct; accordingly, SUBCONTRACTOR grants CONTRACTOR the right to temporarily segregate items, and related paperwork, that are suspected to be S/CI, pending a determination by National Nuclear Security Administration (NNSA) or Department of Energy (DOE) officials whether the segregated items should be impounded as evidence.

- In the event NNSA/DOE directs CONTRACTOR to impound the segregated items, no liability shall be asserted or enforceable against CONTRACTOR, NNSA, or DOE because of the impoundment, all such liability being expressly waived by SUBCONTRACTOR or any person claiming any right or interest under this subcontract in the impounded items.

CONTRACTOR shall incur no liability for failure to return impounded items to SUBCONTRACTOR and does not assume any liability for loss or damage to the items impounded or temporarily segregated pursuant to this clause regardless of the circumstances under which said loss or damage may have occurred, and whether the impounded or temporarily segregated items are in CONTRACTOR’S possession or under its control.

- Nothing in this clause shall limit CONTRACTOR’S right to reject S/CI, and related paperwork, as nonconforming, to deny payment for such items, to return such items to SUBCONTRACTOR once NNSA/DOE has released the items, or to assert other remedies provided under this Subcontract or by law.

- Suspect/Counterfeit Fasteners

LANL will not accept any hex-head cap screws (bolts) with any manufacturer's insignia identified on the Suspect Fastener Headmark List.

- Suspect/Counterfeit Flanges

LANL will not accept any foreign manufactured flange and specifically any flange identified as China A-105 B16.

To access the Suspect/Counterfeit Items Resource Handbook, DOE-HDBK-1221

Section 11 5311.10 - Attachment 4

Inspection Data Sheet -- Glovebox

| INSPECTION DATA SHEET – GLOVEBOX [Identification Number] |
| --- |
| Supplier: | Purchase Order No.: | Date of Inspection: | Inspector: |
| FABRICATION PHASE |
| ( ) Review of QA documentation | ( ) Dimensional inspection weld fit-up |  |
| ( ) Placement of lead sheet prior to installation of stainless steel cladding |
| SHOP ACCEPTANCE TEST |
| ( ) Review of QA documentation | ( ) Inspection of enclosure welds | ( ) Leak Test(s) |
| ( ) Final dimensional inspection | ( ) Inspection of enclosure surface finish |
| SHIPPING AND DELIVERY |
| ( ) Inspection at the delivery site | ( ) Review of QA documentation package | ( ) Final Acceptance |
| Inspection results and comments: |
| ( ) Satisfactory ( ) Unsatisfactory |
| NCR/SDDR Number if Unsatisfactory: |
| Cite the specific QA documentation that was reviewed: |
| INSPECTION APPROVALS | ( ) Approved | ( ) Disapproved | ( ) Approved as noted below |
| INSPECTOR’S SIGNATURE: | DATE: |
| STR APPROVAL: | DATE: |