SECTION 11 5311.12

GLOVEBOX INSTALLATION

*************************************************************************************************************

LANL MASTER SPECIFICATION

Word file at http://engstandards.lanl.gov

This template must be edited for each project.  In doing so, specifier must add job-specific requirements.  Brackets are used in the text to indicate designer choices or locations where text must be supplied by the designer.  Once the choice is made or text supplied, remove the brackets.  The specifications must also be edited to delete specification requirements for processes, items, or designs that are not included in the project -- and specifier’s notes such as these.  To seek a variance from requirements in the specifications that are applicable, contact the Engineering Standards Manual Mechanical POC. Please contact POC with suggestions for improvement as well.

When assembling a specification package, include applicable specifications from all Divisions, especially Division 1, General requirements.

This specification serves as a template.   This specification was prepared by an organization operating under a quality assurance program that meets the requirements of DOE Order 414.1D, ASME NQA-1, and LANL SD330 and is suitable for ML-1 through ML-4 projects.    Implementation of this specification requires modification to the specification to meet project-specific requirements.  Responsibility for application of this specification to meet project-specific requirements lies with the organization modifying or implementing the specification.  The organization modifying the specification shall apply a graded approach to quality assurance based on the management level designation of the project.  When this specification is used with nuclear facilities subject to DOE Order 414.D, modification to this specification must be performed by an individual or organization operating under a quality assurance program that meets the requirements of that DOE Order.

This specification primarily defines requirements for installation of stainless steel gloveboxes used for confinement of nuclear materials within DOE-owned and operated, non-reactor, nuclear facilities at Los Alamos National Laboratory (LANL).

The glovebox installation process covers overall glovebox assembly design, installation design, installation of the glovebox into the facility, installation of glovebox equipment, connection of facility systems to the glovebox, and final inspection and testing.  These activities may be performed by a variety of personnel.  The scope of this specification involves installation of the glovebox into the facility, connection of facility systems, and final inspection and testing.  Activities associated with design for installation and installation of equipment into the glovebox is not presented herein.  If these activities are desired to be performed by the glovebox installation subcontractor, then these scopes of work must be added.

Many of the specification sections below have been developed with reference to standards established by the LANL TA-55 Facility Operations engineering group.  These standards are available for reference by contacting them.

****************************************************************************************************************

GENERAL

summary

Section Includes

Installation requirements to the glovebox installer for nuclear material gloveboxes in LANL facilities.

Reference to other sections of the LANL Master Specifications for more specific requirements, since gloveboxes typically interface with many different engineered systems.

DESCRIPTION

Scope

The technical requirements for the on-site storage, rigging, handling, materials for installation, installation, examination, testing, inspection, and quality assurance (QA) of gloveboxes (GBs), their support stands, and components or appurtenances.

GBs used for the primary and secondary confinement of nuclear materials.  Additional technical requirements are provided in the Subcontract drawings.  Any additional requirements specific to a given GB are identified in Subcontract documentation identified in Division 1 documents and elsewhere.

Acronyms

AGS: American Glovebox Society

ANSI: American National Standards Institute

ASHRAE: American Society for Heating, Refrigerating, and Air Conditioning Engineers

ASME: American Society of Mechanical Engineers

ASNT: American Society of Nondestructive Testing

ASTM: ASTM International (formerly American Society for Testing and Materials)

AWS: American Welding Society

CAM: Continuous air monitor

cfm: cubic feet per minute

CFR: Code of Federal Regulations

CMTR: Certified Material Test Report

CoC: Certificate of Conformance

CWI: Certified Welding Inspector

DOE: Department of Energy

fpm: feet per minute

HEPA: high-efficiency particulate air

LANL: Los Alamos National Laboratory

LVCCW: Limited Volume Circulating Chilled Water

MASS: Material Accountability Safeguards and Security

ML: management level

NDE: nondestructive examination

NEC: National Electrical Code

NFPA: National Fire Protection Association

NPCCW: Negative Pressure Circulating Chilled Water

OSHA: Occupational Safety and Health Administration

PF: Plutonium Facility

PPCCW: Positive-Pressure Circulating Chilled Water

PRV: pressure regulating valve

psi: pounds per square inch

psig: pounds per square inch gauge

QA: quality assurance

QC: quality control

TA: Technical Area

SMACNA: Sheet Metal and Air Conditioning Contractor’s National Association

SOP: Standard Operating Procedure

SS: stainless steel

STD: standard

w.c.: water column

References

Codes, specifications, and standards referred to by number or title form a part of this specification; they include the following references.  Use the latest revision at the time of codes, specifications, and standards referenced below and within this specification award of subcontract, unless otherwise stated.

10 CFR 830, Nuclear Safety Management

29 CFR 1910, Occupational Safety and Health Administration (OSHA) Standards

29 CFR 1910.253, Oxygen Fuel Gas Welding and Cutting

29 CFR 1910.254, Arc Welding and Cutting

American Conference of Governmental Industrial Hygienists (ACGIH),  Industrial Ventilation, A Manual of Recommended Practice

AGS-G001, Guideline for Gloveboxes

AGS-G005, Standard of Practice for the Specifications of  Gloves for Gloveboxes

AGS-G006, Standard of Practice for the Design and Fabrication of Nuclear-Application Gloveboxes

AGS-G010, Standard of Practice for Glovebox Fire Protection

ASME B30.20, Below the Hook Lifting Devices

ASME B31.3, Process Piping

ASME B31.9, Building Services Piping

ASME Boiler and Pressure Vessel Code (B&PVC), Section IX: Qualification Standard for Welding and Brazing Procedures, Welders, Brazers, and Welding and Brazing Operators Welding and Brazing Qualification

ASME NQA-1, Quality Assurance  Requirements for Nuclear Facility Applications

ASNT Recommended Practice No. SNT-TC-1A, Personnel Qualification and Certification in Nondestructive Testing

ASTM A501, Standard Specification for Hot-Formed Welded and Seamless Carbon Steel Structural Tubing

ASTM A36, Standard Specification for Carbon Structural Steel

ASTM A193, Standard Specification for Alloy-Steel and Stainless Steel Bolting for High Temperature or High Pressure Service and Other Special Purpose Applications

ASTM A269, Standard Specification for Seamless and Welded Austenitic Stainless Steel Tubing for General Service

ASTM A312, Standard Specification for Seamless, Welded, and Heavily Cold Worked Austenitic Stainless Steel Pipes

ASTM E 498, Standard Practice for Leaks Using the Mass Spectrometer Leak Detector or Residual Gas Analyzer in the Tracer Probe Mode

ASTM E 499, Standard Practice for Leaks Using the Mass Spectrometer Leak Detector in the Detector Probe Mode

ASTM A325, 	Standard Specification for Structural Bolts, Steel, Heat Treated, 120/105 ksi Minimum Tensile Strength

Withdrawn ASTM A1011, Standard Specification for Steel, Sheet and Strip, Hot-Rolled, Carbon, Structural, High- Strength Low-Alloy, High-Strength Low-Alloy with Improved Formability, and Ultra-High Strength

AWS D1.1, Structural Welding Code –Steel

AWS D1.6, Structural Welding Code Stainless Steel

AWS QC 1, Standard for AWS Certification of Welding Inspectors

DOE O 420.1C, Facility Safety

DOE O 430.1B Chg. 2, Real Property and Asset Management

DOE-STD-1020, Natural Phenomena Hazards Analysis and Design Criteria for Department of Energy Facilities

DOE-STD-1090, Hoisting and Rigging

DOE-STD-3020, Specification for HEPA Filters Used by DOE Contractors

LANL Engineering Standards Manual (ESM), STD-342-100

LANL SD 100, Integrated Safety Management System Description Document with embedded 10 CFR 851 Worker Safety and Health Program

LANL SD 130, Nuclear Criticality Program

LANL P 121, Radiation Protection

LANL P 101-4, Forklifts and Powered Industrial Trucks

LANL P 101-6, Personal Protective Equipment

NFPA 70, National Electrical Code

NFPA 72, National Fire Alarm and Signaling Code

NFPA 86, Standard for Ovens and Furnaces

NFPA 101, Life Safety Code

NFPA 801, Standard for Fire Protection for Facilities Handling Radioactive Materials

SMACNA Standards

Underwriters Laboratories (UL) Standards

related sections

Section 01 1116, Work by Owner

Section 01 6000, Product Requirements

Section 01 2500, Substitution Procedures

Section 01 4455, Onsite Welding and Joining Requirements

Section 01 7700, Closeout Procedures

Section 01 7839, Project Record Documents

Section 05 0521, Post-Installed Concrete Anchors - High Confidence

Section 11 5311.08, Glovebox Design

Section 11 5311.10, Glovebox Fabrication

Section 11 5311.14, Glovebox Gloves

Section 11 5311.16, Glovebox Feedthroughs, Hermetically-Sealed

Section 11 5311.17, Glovebox Instrumentation

Section 11 5311.18, Glovebox Atmosphere Regenerable Purification Systems

Section 13 4800, Sound, Vibration and Seismic Control

Section 22 0529, Hangers and Supports for Plumbing and Equipment

Section 22 0554, Identification for Plumbing, HVAC, and Fire Piping and Equipment

Section 23 4133, High Efficiency Particulate Filtration

Section 23 3225, Bag-in Bag-out Housing

Section 23 0593, Testing, Adjusting and Balancing for HVAC

Section 26 0519, Low Voltage Electrical Power Conductors and Cables

Section 26 0526,  Grounding and Bonding for Electrical Systems

Section 26 0533, Raceway and Boxes for Electrical Systems

Section 26 5100, Interior Lighting

Section 40 0511, Compression Fittings on Copper and Stainless Steel Tubing

Section 40 0513, Process Piping

hoisting and rigging

Perform onsite hoisting and rigging of GBs in accordance with LANL P101-4, “Forklifts and Powered Industrial Trucks” and DOE-STD-1090, “Hoisting and Rigging.”

Rig and hoist GBs in a manner to prevent temporary or permanent distortion of the GB shell (distortion can affect the leak integrity of the GB).

prerequisites

Prior to performing work, plan the installation process and perform work in accordance with LANL’s Integrated Safety Management System as described in LANL SD 100, Integrated Safety Management System Description Document with embedded 10 CFR 851 Worker Safety and Health Program

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall determine whether a Radiation Work Permit will be necessary to install the GB.  An RWP may not be necessary for installation of a GB into a facility where radiation work was previously not performed.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Prior to performing work, develop and establish approval of a Radiation Work Permit (RWP).

Review drawings, details, manuals, and other materials required for installation of GBs prior to performing installation work.

When working with potentially contaminated systems, ensure catch bags, filters, or other mechanisms are in place to contain contaminants prior to performing work.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall develop a list of submittals for each glovebox project since submittals will be project-specific.  Submittals are intended to verify that appropriate materials are used and to report progress and actions performed during installation to the appropriate authority.  The reviewing and approving authority for submittals may vary from facility to facility and from project to project.  The author needs to define the routing of submittals from the submittal generator to the receiving official.  The timing of submittals, numbers of copies, and whether review and/or approval is needed must be defined by the author.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

submittals

Provide reference to LANL Subcontract Number, Glovebox Number, Glovebox Title, and Drawing Number on correspondence.

Provide submittals in accordance with Project requirements.

For all materials used, provide Certified Material Test Report (CMTRs) or (Certificates of Conformance (CoC) if unavailable from manufacturer) indicating compliance with required chemical and physical properties and the test(s) performed to the applicable nationally recognized standards.

Provide other submittals detailed in Quality Assurance below and throughout this Section.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The following quality assurance requirements are consistent with installation of an ML-2 GB.  Where other quality assurance and quality control requirements are needed, modify the following section accordingly.  For instance, the supplier may also apply a QA program in accordance with basic requirements of 10 CFR 830.122.  Add requirements for QA Programs compliant with 10 CFR 830.122 to the specification as necessary.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

quality assurance / quality control

    - The installer shall implement and maintain a quality assurance program in accordance with 10 CFR 830 Quality Assurance; DOE Order 414.1D, Quality Assurance; and  Part I of ASME NQA-1-2008 with ASME NQA-1A-2009 addenda, QA Requirements for Nuclear Facility Applications, and the following additional requirements:

      - Suspect/counterfeit items controls and processes shall be in accordance with DOE Order 414.1D, Attachment 3, Suspect/Counterfeit Items Prevention.

      - Commercial Grade Dedication shall be performed in accordance with ASME NQA-1-2008/-1A-2009, Part II, Subpart 2.14, Quality Assurance Requirements for Commercial Grade Items and Services.  Prior to fabrication, the supplier shall submit for concurrence dedication plans or procedures including the essential elements of the dedication process. Prior to shipment, the supplier shall submit other dedication records for concurrence including the following:

Commercial grade item or service procurement documents,

Technical evaluations

Critical characteristics identification and acceptance criteria

Test reports or results, inspection reports, analyses reports

Commercial grade survey reports

Source verification reports

Historical performance information

Dedication report containing sufficient data to accept the item or service.

Address the following in the installer’s QA Manual (it is acceptable to reference the procedures in the following sections in the manual):

Installation Quality Control (QC) Procedures: Provide list of procedures to be followed and submit per Project submittal requirements.  LANL in cases may waive submittal, where procedures have previously been evaluated.  Maintain a list of quality procedures, including the revision number or date of approval.

Personnel Certifications: Ensure that supplier personnel assigned to GB installation including welding per ASME B&PV Section IX, assembly, testing per ASNT SNT-TC-1A, and inspections are fully qualified to perform their respective job functions.  Submit required personnel certifications per Project submittal requirements.

Test Reports: Ensure that tests performed in support of the GB installation, welding, assembly, testing, and inspection are fully documented.  Submit test reports per Project submittal requirements.

Material Certifications: Provide material certifications for installation materials including legible copies of mill test reports indicating chemical analysis, physical test data, and heat number.

As-Built Drawings: Submit as-built drawings to reflect modifications or deviations to the subcontract drawings during installation.

QA Document Package: Submit documents identified in this specification as a part of the QA Document Package.  Complete three bound or stapled document packages containing these documents required with Project submittal requirements.  Torque Maps: Submit Torque Map identifying location of fastener, required torque, applied torque, and calibration data of torque wrench used to tighten fastener.  Supply Torque Map for fasteners on windows, service panels, access panels, support stand anchors, support stand fasteners, and any other fastener with a specified torque in this specification.

Weld Map: Submit Weld Map identifying location of weld and heat number of materials being welded for all welds to glovebox shell or appurtenances.

Quality assurance records specific to GB installation shall furnish evidence that items or activities meet identified quality requirements.

- Records shall be legible.

- Records shall be traceable to associated items and activities and accurately reflect the work accomplished or information required.

- Records to be generated, supplied, or maintained shall be specified in applicable documents, such as design specifications, procurement documents, test procedures, and operational procedures.

products

product options and substitutions

Comply with Section 01 2500, Substitution Procedures.

materials

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The design agency shall size anchor bolts for their intended purpose.  Anchors should be sized along with the seismic calculations performed in accordance with Section 11 5311.10.  Anchor sizing shall include material selection, embedded length selection and load calculations.  The following anchor specifications are for typical anchors used at LANL.  Calculate specific anchor sizes for each project specific application.  Reference ESM Ch 5 and Section 05 0521.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Anchors

Stainless steel, undercut-type, [1/4-, 3/8-, 1/2-, 5/8- or 3/4-inch diameter] anchors with [X-inch] embedment length as manufactured by [Maxi-Bolt or Hilti], meeting the chemical and physical requirements of ASTM A193, ASTM A325 and ASTM A513.  [Size and select anchors in accordance with manufacturer’s instructions and ESM Structural Chapter.]

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Materials meeting ASTM specifications may be substituted with materials meeting ASME Boiler and Pressure Vessel Code, Section II, Part C, SA-series specifications.  For instance, SA-182 may be substituted for ASTM A182.
((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Utility Penetration Plugs

Stainless steel of same grade as the glovebox shell material, type 304 or type 316 and that meet the physical requirements of ASME B16.11.

Plugs shall also meet the chemical and physical requirements of ASTM A182 and ASTM A403 and be dual marked 304/304L or 316/316L accordingly.

Do not use cast material for utility penetration plugs.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

High-density polyethylene (HDPE) provides very good chemical resistance.  Pressure rating: 150 psi at 77°F (25°C). Temp range: –65 to 190°F (–53 to 87°C).

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

For highly corrosive glovebox environments: high-density polyethylene (HDPE) utility plugs in lieu of stainless steel plugs on the interior of the glovebox only.

Use Teflon tape and Tru-Blu sealant in conjunction with threaded utility plugs and caps.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

GB primary supply and return confinement plenums consist of gloveboxes and conveyor tunnels which are served by a 100% once-thru exhaust system consisting of ductwork, HEPA filters, GB type plenums and fans.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Duct

GB Supply & Return Ducted Plenums

Material of Construction

For GB supply and exhaust air ducts originating at the conveyor tunnel or gloveboxes and terminating in the exhaust airway: stainless steel pipe or tube ductwork per the design drawings

Stainless steel pipe/tubing: Type 304L, circular duct with a minimum 16 gauge, with seamless or welded leak tight joints..

        - Provide supply and exhaust plenums that are completely sealed and leak tested.  Prior to final leak testing, the internals such as cooling coils, sprinklers, fire protection and all the associated piping, controls, sensing devices, electric wiring and fixtures are to be installed and connected.

        - Provide a leak test of the duct system to ensure the air tightness of all connections showing no measurable leak at 0.5 psig positive or negative pressure.

        - Employ additional weld examination criteria (i.e., visual, light level of at least 100 foot-candles and liquid dye penetrant in accordance with Appendix VIII, Section VIII of the ASME Boiler Pressure Code.) on both seal and strength welds.  Complete all welding of internals to skin and penetrations of plenums before the final overall leak test.

Provide all duct construction, minimum material gauges, longitudinal joints, transverse joints, and reinforcing of sheet metal ductwork in accordance with ASHRAE and SMACNA standards.

Construct all ducts, other than glovebox supply and return plenum, from black carbon steel, galvanized sheet, or stainless steel based on the temperature and reactivity of the gases being conveyed.

When the gases being conveyed are corrosive, use duct linings or coatings to extend duct service life as required.

Use round ducts unless prohibited by physical constraints.

Ensure all materials meet or exceed design specifications and are free from damage or manufacturing defects.

Material specifications for all ducts other than glovebox supply and return plenums:

Hot-rolled steel sheet in accordance with ASTM A1011.

Hot-rolled steel plate in accordance with ASTM A36.

Galvanized sheet in accordance with ASTM A527.

Stainless steel in accordance with ASTM A167 or A269 for tubing.

Piping

- Refer to Section 40 0513, Process Piping

Wiring

Refer to Section 26 0519, Low Voltage Electrical Power Conductors and Cables.

Conduit

Refer to Section 26 0533, Raceway and Boxes for Electrical Systems.

Receptacles

Refer to Section 26 2726, Wiring Devices.

Wireways

Refer to Section 26 0533, Raceway and Boxes for Electrical Systems.

Instruments

Refer to Section 11 5311.17, Glovebox Instrumentation.

Light Fixtures

Locate general lighting exterior to GBs (low heat), mounted above or near a window section built into the GB shell.  Install enclosed LED type specifically modified for GB use. Power interior lighting from dedicated duplex outlets above the GB.  Ensure the lights chosen will not create a fire risk by overheating and are noncombustible construction.

Provide local directional LED type task lighting.  Secure task lighting to the GB.  Power interior lighting from duplex outlets above or below the GB.  Ensure task lights chosen will not create a fire risk by overheating the exposed GB surface.

Install LED fixtures to provide lighting levels of approximately 100 foot-candles at work surface.  Provide methods of adjusting lighting fixtures to minimize glare.

Install luminaries with baffles to diffuse light and ensure light tube is not directly visible to a user’s eye.

Install LED task lights where space is limited in bulkhead-type fixture.  Avoid close proximity to flexible glovebags because melting or fire could cause breach of containment.

Avoid high intensity light fixtures near GB windows, they have been known to break GB window due to excessive heating.

For torque requirements, refer to Section 3.2.C.3 Window Installation Fasteners

Carbon steel and stainless steel fasteners as specified in Section 11 5311.10, Glovebox Fabrication.  Prevent the use of counterfeit or suspect fasteners by following the guidance in Section 11 5311.10. Fasteners not marked in accordance with the ASTM Standards, shall be considered nonconforming and subject to rejection.

Provide Certified Material Test Reports (CMTR) for fasteners to LANL for review and approval.  With prior LANL approval, provide Certificates of Conformance (CoC) in lieu of CMTR for fasteners.  Review and approval of submittals by qualified engineering authority.

Use UNC-series threads for fasteners and components in accordance with ASME B18.2.1.

Welding Materials

Welding materials including filler rod meeting the requirements of Section 11 5311.10, Gloveboxes and ESM Chapter 13, Welding and Joining.

Provide Certified Material Test Reports (CMTR) for weld filler to LANL for review and approval.

Weld Studs – Ensure capacitive discharge (CD) weld studs welded to the interior and exterior of the glovebox are the same type stainless steel as the glovebox shell. Provide CMTR with weld studs.

Electrical Feedthroughs

Refer to Section 11 5311.16, Glovebox Feedthroughs Hermetically-Sealed.

Glovebox Atmosphere Regenerable Purification Systems (Dri-Trains)

Refer to Section 11 5311.18, Glovebox Atmosphere Regenerable Purification Systems.

Glovebox Gloves

Refer to Section 11 5311.14, Glovebox Gloves.

execution

structural

Install glovebox support stands in accordance with the project design documents.

Make surveys of the laboratory floor to determine the location of reinforcing bar imbedded in the concrete laboratory floor (e.g., ground-penetrating radar – GPR).  Do not drill anchor holes that intersect with the reinforcing rod.

Torque bolts per project documents. Document torque applied to fasteners on a Torque Map of the GB.  Indicate the location of fastener, torque applied, and calibration data for the torque wrench used to tighten fastener.  Submit Torque Map with project records.

Properly secure heavy objects on top of or inside GB.

glovebox integrity

Welding and Surface Finish

Maintain surface finish on gloveboxes in accordance with the requirements of LANL Master Specifications Section 11 5311.10, Glovebox Fabrication.  Maintain surface finish requirements in accordance with requirements defined on drawing 26Y-202001.

Grind and polish all welds between the glovebox and exterior or interior appurtenances attached during installation.  Perform all grinding and polishing operations in accordance with Section 11 5311.10.

Surfaces shall be free of weld spatter, burrs, sharp edges, corners, loose stranded wires and projecting pins.

Polish damaged surfaces to a 32 micro-inch, arithmetic average, finish.

Provide a Weld Map to document all welds applied to the glovebox during installation.  Document the location of the weld on the glovebox and the heat number of material being welded.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall determine and specify the type of gloves to be installed in the glovebox, depending upon requirements for resistance to chemicals, radiation shielding, and ergonomics.  The type of glovebox and the process operation conducted inside the glovebox determines the type of glove used.  In such areas as plutonium-238 labs, a more detailed procedure for installation is required, than specified below, due to the high specific activity of plutonium-238.  The following procedure for glove installation is one used by TA-55 PF-4 and RLUOB Facilities.  The author shall use this procedure unless other facility-specific procedures for glove installation exist at the site of glovebox installation.

Reference the American Glovebox Society, AGS-G005, “Standard of Practice for the Design and Fabrication of Gloves and Transfer Sleeves” and the LANL Construction Specification, Section 11 5311.10, “Glovebox Fabrication,” for information pertaining to glove selection and use.

Consider that in some cases, where beta emitters are being handled, unshielded glovebox gloves will reduce the extremity dose rate.  Consult LA-UR-99-3596, Photon and Electron Shielding for Neptunium-237, for further information.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Gloves

Refer to Section 11 5311.14, Glovebox Gloves.

Verify no nicks are around gloveports before installing the gloves.

Install gloves or gloveport plugs in all gloveports provided on the glovebox shell.

Gloveport covers shall be installed on each of the glove openings that are used infrequently (gloveports not used within 60 days) in accordance with AGS-G010. The covers shall be constructed of non-combustible material and have the capability of positioning the gloves outside the glovebox.

Inspect gloves prior to use for cracking, tears, punctures, and burns.  Perform inspection in accordance with the facility approved glove inspection procedure.

Perform the following procedure for installing gloves in gloveports manufactured by Central Research Laboratories (CRL).

Place the bead of the glove into the narrow groove of the plastic support ring.

Apply a light coating of vacuum grease to the O-ring.  Install the O-ring over the glove and onto the wide groove of the support ring.  Make sure that the glove is not wrinkled under the O-ring and that the bead is still in the narrow groove.  This unit will be referred to as the glove-support ring subassembly.

Retract the inner housing of the ejection tool by turning the hand wheel counterclockwise.

Retract the plunger and remove the swaging collar from the ejection tool by rotating it in either direction until it releases from the bayonets.

Apply a very light coating of vacuum grease periodically to the inside surface of the swaging collar.

Install the glove-support ring subassembly into the swaging collar, orienting the thumb of the glove with the plunger notch on the swaging collar.  This will result in a “thumb-up” position of the glove when installed.

Engage the swaging collar to the bayonets on the ejection tool and rotate until the plunger locks into the notch of the swaging collar

Check the thumb of the glove and the plunger for proper orientation.

Insert the glove into the inner housing.

Orient the bayonet of the swaging collar (the bayonet that is in line with the plunger) with the bayonet opening in the enclosure ring at approximately a 10 o’clock position and rotate the ejection tool clockwise to the stop pin.  Ensure that the notch and plunger on the ejection tool is positioned at 12 o’clock when the tool is properly installed on the GB ring.

By working through the spokes of the ejection tool, apply a light pull force on the glove.  Turn the hand wheel of the ejection tool clockwise until the hand wheel bottoms on the outer housing.

Retract the inner housing of the ejection tool with the hand wheel and remove the ejection tool from the GB ring.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The following procedures for window installation are used by TA-55 PF-4 and RLUOB Facilities.  The author shall use these procedures unless other facility-specific procedures for window installation exist at the site of glovebox installation.  The procedures below describe replacement of windows on gloveboxes.  These procedures should be modified appropriately to describe installation of windows on new gloveboxes when new gloveboxes are being installed.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Windows

Install windows provided by the glovebox supplier on the glovebox in accordance with the following requirements.  Reference Section 11 5311.10 for specifications associated with glovebox windows.

Preparation for Window Installations

Obtain proper windows and gaskets

Obtain and prepare tools and equipment to perform the job

Obtain proper authorization [i.e., Radiation Work Permit (RWP), fixed head readings, nose swipes, and/or survey maps]

Inspect the new window glass pane for proper size and shape, (proper radii is essential for proper seal and fit)

Prepare area for proper contamination control

Grease the grooves of the new gasket and attach it to the new window.  For bolt-on windows, grease the new gasket and attach it to the new window.

Installation of a Bolt-On Window

Install the new window, with its greased gasket attached.

Apply anti-seize to the bolts prior to installation of the frames.

Using an alternating pattern, tighten the screws with either a standard wrench or torque wrench enough to stop leakage around the gasket.  When using a torque wrench, apply the recommended torque of 25 inch-pounds +/- 5 inch-pounds to the screws.  The objective is to tighten the bolts enough to stop leakage around the gasket.

If cracks are evident or produced.  Inspect window penetration and work to flatness tolerance.  Refer to Subsection 3.2.A above.

Document torque applied to fasteners on a Torque Map of the glovebox.  Indicate the location of fastener, torque applied, and calibration data for torque wrench used to tighten fastener.  Submit Torque Map for project record documents. Inspect the window glass for cracks prior to use.

Installation of a Zippered Window

Apply vacuum grease to the outer groove of the zipper window gasket.

Install the greased groove of the new gasket into the window opening.

Insert the new window into the center of the greased groove of the gasket and slide the glass toward one side of the gasket.  When the glass reaches the end of the gasket, use the plastic window tool to finish inserting the glass into the gasket.

Install the zipper seal.

Inspect the window glass for cracks prior to use.

Service and Access Panels

Plug unused penetrations in service panels with a plug as defined in the materials section of this specification. It is good practice to plug both the outside of the service panel and the inside.

Install hermetically-sealed electrical feedthroughs in openings in service panels in accordance with Section 11 5311.16, Hermetically Sealed Glovebox Feedthroughs.

Attach pipe and tubing to female threaded couplings on service panels in accordance with requirements set forth in ASME B31.3.  Apply Teflon tape and Tru-Blu sealant to threads of male pipe couplings prior to installation.

Attach service and access panels to glovebox using gaskets provided with the glovebox from the glovebox manufacturer.

Utilize high-crown acorn nuts, supplied with glovebox from glovebox manufacturer, for attachment of service and access panels.

Torque service and access panel nuts to 25-inch pounds +/- 5 inch pounds.

Document torque applied to fasteners on a Torque Map of the glovebox.  Indicate the location of fastener, torque applied, and calibration data for torque wrench used to tighten fastener.  Submit Torque Map for review and approval.

Use caution; over-tightening can crack the window or cause failure of the weld stud or screw.

Perform post-modification helium leak test of the glovebox including installed service panels and hermetically-sealed feedthroughs in accordance with requirements defined in Section 3.12.

Gaskets

Install gaskets with all service panels, access panels, and glovebox primary confinement penetrations in accordance with project documents.  Gaskets shall be noncombustible material or protected from potential fire exposure by coving/sandwiching between noncombustible materials.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

HEPA filters are used for a variety of gloveboxes (i.e., air atmosphere, isolated air, inert gas flush, high purity inert atmosphere, etc.).  The exhaust for most gloveboxes is through 8-inch or 12-inch diameter HEPA filters.  The glovebox atmosphere is drawn from the glovebox through a HEPA filter located on top of each glovebox mitigating contamination of the ventilation system.  HEPA filters are located on gloveboxes in accordance with DOE Order DOE O 420.1C.  The author shall determine the appropriate size of HEPA filter and housing for a particular process.  Discussion here will be reserved for standard 8-inch or 12-inch diameter primary exhaust filters for most gloveboxes as detailed in LANL TA-55 Drawing Nos. 26Y-202057, 26Y-202059 and 26Y-202130.  It is recommended that each facility develop and implement its own filter preferences and filter changing procedure.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

HEPA Filters

Refer to Section 43 3113, High Efficiency Gas Purification Filters.

Install HEPA filters that meet nuclear facility grade requirements as described in DOE-STD-3020 and the LANL Master Specification Section 23 4133, High-Efficiency Particulate Filtration.

Qualify all personnel who will change exhaust HEPA filters in accordance with facility specific procedures.

Install the HEPA filter sprinkler water discharge cap assembly, if applicable.

Shaft Penetrations

Install shaft penetrations that require rotational or translational mechanisms to penetrate the glovebox so that equipment is integrated into the glovebox.

Align the motor and shaft assembly to prevent excessive vibration of the glovebox or distortion of the glovebox.  Excessive vibration can cause glovebox window breakage and may affect sensitive analytical balances.

Equip static (bolted support housing) and dynamic (rotational shaft) assemblies of the penetration with high integrity seals (i.e., O-ring, hermetic, ferrofluidic, etc.) to maintain confinement integrity of the glovebox.

Include primary, secondary and tertiary seal confinement at both the static and dynamic assemblies at seal installation.  Include a gasket or O-ring seal at the support plate.  Refer to a standard model below that illustrates the seal requirement.

Penetration Wall Principle:

Subject the shaft penetration, with the equipment in place and operational, to a post-modification helium leak test along with the glovebox in accordance with the requirements identified in the testing portion of this specification.

Hoods

Refer to Section 23 3816, Fume Hoods.

mechanical services (all gloveboxes)

Connection to Existing Equipment

Airlocks

          - Moving items into and out of an airlock requires operation of airlock doors that could pinch operator’s hands/fingers and glovebox gloves. Movement of material through an airlock is a mechanical operation involving opening and closing doors. Thus, the potential of getting a hand or a glove caught in the door (or door mechanism) exists.  Ensure adequate clearance between the gloveport nearest the sliding door.

          - Install airlock doors with pneumatic actuators that are structurally stable.  Trim length of door as required to ensure proper travel of pneumatic cylinder shaft.

          - Ensure there are no stops in the door guides that restrict or bind opening of the door.

          - Install airlock spool pieces of sufficient length to install glovebox and accommodate connection of utility lines to side service panel.

          - Ensure gas supplies entering airlock have pressure-regulating valves, check valves gauges near workstations and filters.

          - Set all pressure-regulating valves (PRV) for gas services except instrument air at less than 25 psig.

          - Label all airlock service lines that are traced to facility services headers.

          - Install air-operated cylinders with PRVs and flow-regulating valves.

          - Install air-operated (pneumatic) switches for glovebox doors, mounted approximately 10-in. below the top of the glovebox legs.

          - Install appropriate gasket material and shape for airlock connecting ring assembly to provide adequate seal.  Mount connector rings at 90( ( 3( to axis of enclosure.

          - Perform ground penetrating radar (GPR) to locate concrete structural members prior to setting anchor bolts.

          - Locate approximate location of glovebox structural support plates and approximate leg support heights.  Cut each support leg to individually measured distance, or adjust the legs to the appropriate height.

          - Install legs to complete support structure, weld base plates.

          - Approximate and align glovebox to airlock in the longitudinal and transverse directions.

          - Drill holes at designated location to anchor support structure.

          - Level box by adjusting legs to the appropriate height

          - Secure and seal door to adjacent gloveboxes or dropboxes before breaking seal to airlock. Contact Radiation Control to install signage on the active drop box, once the door is sealed and ensure ventilations is connected.

          - Acquire standard glovebox parts for new spool piece (airlock) with new gaskets (metal or neoprene) and clamps

          - Finalize alignment of gloveboxes with lifting jacks or table.  Verify load capacities of lifting devices to appropriate rated loads and safety factors.

          - Back glovebox to airlock and install gaskets and standard ring assembly.

          - Clamp and torque standard ring assembly “snug tight” or 60 in-lbs. minimum.

          - Document torque applied to fasteners on a Torque Map of the glovebox.  Indicate the location of fastener, torque applied, and calibration data for torque wrench used to tighten fastener.  Submit Torque Map for review and approval.

          - Perform post-modification helium leak test on the airlock in accordance with the testing section this specification.

Open-Front Gloveboxes

Open-Front Gloveboxes (i.e., transfer boxes, introductory boxes) are used to insert or remove materials and/or equipment into the glovebox line.  Inlet air for these gloveboxes is provided from the laboratory rooms.  The exhaust air from these gloveboxes passes through a HEPA filter and is drawn into a recirculation system plenum.

Open-front gloveboxes use a 24-in. square stainless steel HEPA filter.  These boxes need a large capacity filter to handle the volume of air drawn in to ensure the required air flow when the door is open.  Refer to Section 23 4133 for HEPA filter requirements for use on open-front gloveboxes.

For chemical hoods with no radioactive material, the glovebox ventilation system must provide 125-fpm-minimum airflow through full open doors.

For chemical hoods where radioactive material is available, the glovebox ventilation system must provide for 150 fpm through the normal opening.

Open-front gloveboxes must be tested, display a current performance acceptance (yellow) sticker and comply with any conditions stated on the certification.

Before introducing items through hoods, consult with a radiological control technician to determine the radiological controls necessary for the introduction, including, but not limited to, types of PPE and the necessity of radiological control technician coverage.

Conveyors

Refer to Section 41 2225 – Hoists and Trolleys for requirements pertinent to hoists and trolleys used inside GBs.

Install push-button control stations located at each drop box station and at the control console.  The numbered push buttons shall match the numbered drop boxes serviced by the trolley, and each button will light up when the trolley is at the corresponding drop box. The trolley will not move until the hoist is in the UP position and locked in place.

Install a trolley bucket that can be lowered into any drop box in an individual system.

Typical hoist and trolley shall have a load limit of 500 pounds (227 kilograms) with the trolley bucket removed. The load limit for the trolley with the bucket attached is 200 pounds (91 kg) in addition to the weight of the bucket.  This will prevent glovebox containment breach at the system interface.

Install lifting equipment meeting the requirements of ASME B30.20 and the following requirements:

Metal tag or placard displaying the load limits, proof load, date of last proof test and retest interval.

Positive mechanical locking device to prevent inadvertent lowering of the load in the event of lifting mechanism failure.

Permanent mechanical stops.  This will prevent a trolley from falling off the end of a monorail or contacting the interior of the glovebox.

Install independent mechanical and electrical braking systems for electrically powered lifting mechanisms.

Install braking systems capable of braking and safely holding a minimum of 150 percent of rated load.

Use hooks for lifting equipment, including hooks used on slings and cables, with positive safety latching devices across the hook opening.

Attachment points for monorails are typically spaced 8-in. to 12-in. apart and utilize studs to support welded structural plates or angles.

Provide guards made from Lexan or other suitable materials to protect windows.

Cable hoists are preferred over chain hoists.

To prevent glove tears, grind smooth any sharp points, corners, and edges.

Ensure that that objects requiring movement inside gloveboxes designed with L-shaped floor plans and provided with a curved monorail do not contact the inside corner of the glovebox.

Subject conveyor tunnels enclosures to a post-modification pressure test.  No measurable leaks at 0.5 psig positive and negative.

Torque for tie-down requirements are defined in project documents.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

For additional information the author shall reference the American Glovebox Society latest edition, “Standard of Practice for the Design and Fabrication of Gloveboxes for the Containment of Materials that Emit Low-Penetrating Ionizing Radiation,” AGS-G006, and LANL Facility Master Specification Sections 11 5311.08 – Glovebox Design and 11 5311.10 – Glovebox Fabrication.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Gloveboxes and Dropboxes

NOTE: A glovebox system includes the gloveboxes, drop boxes, conveyer tunnels, trunk lines, open front boxes/transfer boxes and fume hoods.  Drop boxes and trunk lines typically interconnect gloveboxes so that hazardous and radioactive material and equipment may be transferred from one glovebox to another.

Completed glovebox installations, including service tie-ins, support systems functions, etc., shall be certified by a subject matter expert (system engineer) according to an established checklist (Attachment 1).  Certification must be approved prior to final connection to the glovebox ventilation system and startup of glovebox operations.

Architectural Requirements

The face of the glovebox is typically aligned with the face of the drop box.

Provide a minimum of 3-in to 6-in. space between the backs of GBs (i.e., a GB line in the center of a laboratory) or between the GB and laboratory wall.

Install GB and use standard gloveport height of 48-inches to 52-inches from laboratory floor to centerline of round gloveport. Use standard glovebox port spacing of 46-inches to 54-inches for oval gloveports.

Provide a rubber mat on the floor and a portable work platform of varying height, at front of the GB.

Ensure adequate vertical spacing to install the GB connection to the ventilation as designed.

Ensure adequate room to access the glovebox filter housing for filter replacement.

Locate gloveboxes and/or glovebox equipment per the drawing details.

Sprinkler protection or the creation a non-accessible, noncombustible concealed space (noncombustible skirting) is required to be installed beneath GBs, where its width and depth both exceed 4-feet and the bottom of the GB is greater than 24 in. above the floor.

Mechanical

Install gloveboxes and hoods to achieve the following face velocities.

- For chemical hoods with no radioactive material, the glovebox ventilation system must provide 125-fpm minimum airflow through full open doors.

- For chemical hoods where radioactive material is available, the glovebox ventilation system must provide for 150 fpm through the normal opening.

- Setup ventilation systems on gloveboxes so that there is a flow of 125 fpm ± 25 fpm through the largest credible breach (gloveport or bagport).  Set-up glovebox normal operating flow, inlet resistance (damper setting), and glovebox pressure  to achieve 125 fpm ± 25 fpm face velocity when the breach occurs, without having to actuate dampers (except in closed loop inert gas systems).

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Author shall reference the Glovebox Standard details from the 26Y-202000 drawing series to refer to design details (i.e., Counterbalance Doors, Round and Square Airlock Assemblies, Exhaust Filter Assemblies, etc.) for standard design and installation criteria.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Install Unistrut welded to the back legs to hold pipe and electric headers flush with the side of the legs (i.e., not extend to the rear beyond the legs).

Install piping from nipples on the front of boxes that have hydrogenous shielding; it should be as close to the box surface and stand as possible.

Install flat gaskets for bolted covers and service panels of black neoprene, 1/8-inch to ¼-inch thick, 40 to 60 durometer, (Shore A) as shown on project drawings

Torque window frame, service panel, or plate cover bolts or nuts to 25 ( 5 in.-lb.

Document torque applied to fasteners on a Torque Map of the glovebox.  Indicate the location of fastener, torque applied, and calibration data for torque wrench used to tighten fastener.  Submit Torque Map for review and approval.

HEPA Filters must be leak tested. One method uses di (2-ethylhexyl) sebacate (called DOS testing), which replaced di-octyl phthalate, or DOP testing.  Refer to Section 23 4133 – Filters for further requirements.

Install GB gas service lines as shown on project documents.

Perform fit-up of compression fittings in accordance with Section 40 0511 of the LANL Master Specifications.

Set pressure regulating valves for gas services except process air at a maximum of 25 psig.

Install industrial water piping into the GB with “deadman” valves to prevent flooding.

Install PPCCW to cool equipment exterior to the glovebox. PPCCW shall not be piped into the GB.

Insulate chilled and heated piping.

After installing wet vacuum piping, wrap all flanges in plastic.

Install supply lines for air driven stirrers, vibrators, pumps, etc., inside gloveboxes with solenoid valves interlocked to glovebox pressure.

Install vacuum pump piping to GBs, exhaust piping, and HEPA filters with the proper ventilation system according to installation drawings.

Install air operated cylinders with PRVs and flow regulating valves.

Install air operated cylinders with shafts penetrating the GB that have HEPA filters and exhaust to the proper ventilating system.

Leak test pressurized piping, vessels, columns, etc., in accordance with LANL Master Specification Section 22 0813, Testing Piping Systems.

Install bent tubing in lieu of elbows unless otherwise indicated.

Install isolation valves for independently controlling liquids and gases to any process system or experiment inside a glovebox.

Welders and brazers shall be certified in accordance with the ASME Boiler and Pressure Vessel Code, Section IX, and ESM Chapter 13, Welding, Joining & NDE.

Install dielectric connections wherever joining dissimilar metals.  Pitch lines in direction of flow.

Install labels on piping in accordance with P&ID drawings, the LANL Engineering Standards Manual, the LANL Master Specifications, and Facility Procedures.

Install Glovebox, piping and ventilation ducting that do not obstruct sprinkler system discharge patterns unless additional sprinklers are installed as required by NFPA 13.

Electrical Requirements

Interlock all glovebox furnaces that have a high vacuum unit attached to cut off the power if vacuum is lost (pressure rise above a set point).  Install a bypass switch to allow over ride of this feature, unless vacuum is credited for fire prevention, then no override is allowed.

Interlock all glovebox furnace overtemperature shut offs to cut off the power if overtemperature condition is reached (temperature rise above a set point).  Require manual intervention to return power to furnace (NFPA 86).

Install one 115V duplex outlet in single section drop boxes.  Install two 115V duplex outlets in double or multiple section drop boxes.

Perform a helium leak test on hermetically-sealed electrical feedthroughs following installation in glovebox penetrations and service panels in accordance with Subsection 3.12 of this specification.

Properly identify circuits at penetrations and breaker panels.

Check voltage and wiring.

Check Photohelic and solenoid wiring.

Ground all GB support stands.

Instrumentation and Control Requirements

Refer to Section 11 5311.17– Glovebox Instrumentation for further requirements.

Install labels to identify instrumentation receptacles and cables.

Install tubing and/or conduit lines that do not go up the front, operating side, or located directly over a GB window or connector ring gasket.

Install indicator gauge dials with an appropriate scale to be visible from the front of the boxes.

Install air operated (pneumatic) switches for GB doors mounted approximately 10-in. below the top of the GB support stand.

Bubblers (pressure relief devices), must be filled with oil to the proper level.  Install bubbler bypass piping with manual and solenoid valves. A Photohelic is required for all bubbler bypass installations.

Install air purification systems (Dri-Trains) with HEPA filters, pressure regulating valves and flow meters (when required) in gas supply lines.  Refer to Section 11 5311.18 – Glovebox Atmosphere Regenerable Purification Systems for further requirements.

“Zero” Magnehelic and Photohelic or other pressure controllers and test for proper operation of makeup valve and exhaust (bubbler bypass) solenoid valve in order to respond to changes in GB negativity.

For glovebox operations (see Para 3.2.2), where the inside of the glovebox has the potential for a hydrogen build-up, install a hydrogen detector.

For GBs exhibiting the potential for a hydrogen environment, install a hydrogen alarm switch that actuates solenoid valves on the hydrogen lines outside the building or in the hydrogen supply cabinet.

Fire Protection Requirements

Refer to Section 11 5311.17 – Glovebox Instrumentation for further fire detector requirements.

Fire detection shall be installed in each GB and drop box, using 8 ft spacing criteria and at more than 4 ft. from the nearest wall or as otherwise directed in the fire hazard evaluation (FHE).

When the glovebox fire detection system includes thermal detectors installed in glovebox wells, 140 °F rated detectors shall be installed unless the glovebox contains a heat sources, then the thermal detector shall be rated at 190 °F.

Install a means to prevent fire spread in all drop boxes.

Fire screens shall be installed on exhaust GB HEPA filter cutouts to the GB on the interior.

Install inlet GB HEPA filter sprinkler water guard to prevent fire sprinkler water from entering the glovebox through the inlet HEPA filter in the event of sprinkler discharge.

Combustible loading inside gloveboxes is limited to the equipment and materials specifically required for the day’s glovebox installation work.  Installation of gloves must ensure that gloves do not contact hot surfaces inside the glovebox. Gloves shall be pulled and secured outside of the GB when not in use.

Install a fire shield to protect zippered windows with neoprene gaskets where the gaskets are exposed to the interior of the GB.

Special Installation Instructions

The final helium leak test shall be performed after all GB utilities have been connected, all penetrations have been installed, and the GB installation in the operating facility is complete. The removal and reinstallation (including rework and repair) of a sealing component from the confinement boundary of the GB will negate the validity of any previously performed helium leak test and shall require a final helium leak test to be performed. The final helium leak test shall be performed prior to connection to the ventilation system.

Environmental Requirements

Reference Part 3, Article 3.4 for instruction on internal environmental parameters.

Connection to Ventilation

Zone 1 or Primary Confinement

Construct ducts, and use expansion joints, connections and seals in accordance with ASHRAE and SMACNA standards.

Zone 2 or Operator Environment

Materials and techniques of construction should be in accordance with ASHRAE and SMACNA standards latest revisions.

General Piping and Tubing Connections

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall develop specifications for connection of glovebox floor drains to facility systems.  Consult with LANL Criticality Safety group and facility authorization basis when developing specifications for this section.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Floor Drains

Connect glovebox floor drains to facility systems in a criticality-safe fashion.  Consult with the Nuclear Criticality Safety and Safety Basis groups prior to connection of glovebox floor drains to facility systems.

mechanical services (inert atmosphere gloveboxes)

Glovebox Atmosphere Regenerable Purification System (Dri-Train) Interface

Refer to Section 11 5311.18 – Glovebox Atmosphere Regenerable Purification System

Complete isolation from adjacent gloveboxes, drop box, and conveyer tunnel and tight controls are necessary to provide for a high-purity inert recirculating gas regeneration system.

Control oxygen concentration in the [2] parts per million range, and nitrogen concentrations in the [30] parts per million range.

NOTE: Purification to remove oxygen and nitrogen is accomplished through use of a Dri-train unit.  The Dri-train unit uses a blower to recirculate the glovebox atmosphere through a dryer/oxygen scavenger (a molecular sieve and copper oxide reactant).  The Dri-train unit is connected to the glovebox by piping.

Service Classification and Description

Use hard copper, ASTM-B88, hard temper, type “L” wall thickness.  Use soft copper for flexible connections including instrument leads from branch block valves to instrument.  (Ref.: Part 2, Subsection 2.2.D., “Piping”)

Operating ranges: Maximum Pressure, 0 to 125 psig, Maximum Temperature, 35 (F to 180 (F

Construction Code: ASME B31.3

Material Standards: ASTM-B88, ASTM-B32, ASTM-B75, ANSI-B16.22, and ANSI/ASME B1.20.3.

Provide for glovebox exhaust through HEPA filter assemblies and a bubbler (pressure relief devices) to regulate pressure with a Photohelic-activated bypass (Ref.: Section C of this section).

Establish bubbler bypass pressure setpoints (i.e., make-up gas will start to flow at –0.7 inch w.c., and it will build up pressure until it reaches –0.4-inch w.c.  At –0.4 inch w.c., the make-up gas switches off and the bypass valve opens, lowering the pressure.  Should a pressure build-up overwhelm the bypass, the bubbler is set to bleed off at –0.2 in w.c.)

Use bent tubing in lieu of elbows for all purified gas and chilled water unless otherwise indicated.  Hard solder all purified gas and chilled water joints unless otherwise indicated.  Butt-weld exhaust vent joints unless otherwise indicated.

Install flex connectors with solder cup ends for vent and purified gas piping at Dri-train unit.

Install quick disconnects “snap tight” with textile-braided neoprene hose for Positive pressure circulated chilled water supply/return at Dri-train unit.

Install the large 1-1/2 inch valves on Dri-Train gas lines as close to the glovebox nipples as possible.  On gloveboxes using hydrogenous material, the shielding only extends slightly above the large windows on the front of the box; therefore, these valves will be operable if placed high on the box sides.

Ensure that all connections for recirculating gas and regeneration are accessible through the back panel.

Ensure that the flow rate of inert gas through the regenerative chambers is adjustable through the range of 0 to 40 cubic feet per minute.

Ensure that the gas entering the return line to the glovebox is not more than 10 °C above ambient room temperature.

Regeneration gas must be nitrogen or argon with 3% to 5% maximum hydrogen content, pressure-regulated to 25 psig.

Install vent lines from the process vent during regeneration (¼-inch NPT) and vacuum pump (1-inch NPT), and connect to the suitable exhaust lines.

Perform helium leak testing of the entire system in accordance with ASTM E499 (tracer probe mode).  Test the system to verify no detectable leak greater than 1 x 10-6 std cc/sec with helium leak-test equipment capable of detecting a smaller leak.

Once-Through System Interface

The GB is isolated from adjacent GBs and an inert gas is flushed through.  Install a bubbler in the glovebox line after the HEPA filter on top of the glovebox to control pressure.

Install a manual control valve to set the atmosphere in the glovebox to the desired composition.  The bypass is closed and the supply flow rate is manually controlled to less than 5 cfm.

Once-through system interfaces maintain less than 5% oxygen (≤ 4% for pyrophoric metals) or water vapor atmosphere environment.

Service classification and description as per Recirculating Gas Regeneration System Interface.

Use bent tubing in lieu of elbows for all purified gas tubing unless otherwise indicated.  Hard solder purified gas joints unless otherwise indicated.  Butt-weld exhaust vent joints unless otherwise indicated.

Install a HEPA filter, check valve, pressure regulator, and flow meter in all once-through interfaces with gloveboxes in gas supply lines.  Refer to Section 23 4133.

Pressure Relief Devices

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall determine the pressure requirements and applicability of utilizing a bubbler as the pressure relief device to control inert atmosphere gloveboxes.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Bubbler Installation

Bubbler units provide glovebox overpressure relief when an inert atmosphere is used, or on any glovebox that might over pressurize.  See illustration below.

Preferred liquids used for balancing glovebox internal pressure with consideration for appropriate density, viscosity and flammability are:

Dow Corning 200

Duo Seal 1407

Unit assembly must be vacuum leak checked for a helium leak rate not to exceed 1 X 10-3 cc atm/sec.

Install the bubbler assembly in the exhaust line of a glovebox.

Install welded ¾-in type 304 SS half couplings to Zone 1 ductwork and downstream HEPA Filter housing to accommodate all SS bypass piping and fittings.  Perform fit-up of compression fittings in accordance with Section 40 0511 of the LANL Construction Specification.

Install bubbler mating flanges to HEPA filter and Zone 1 ductwork with the appropriate clamp and O-ring seal (i.e., Aeroquip Corp. or approved equivalent).  This is relevant to the typical Bubbler illustrated in Figure 2.

Torque clamp wing nut assembly “snug tight.”

Document torque applied to fasteners on a Torque Map of the glovebox.  Indicate the location of fastener, torque applied, and calibration data for torque wrench used to tighten fastener.  Submit Torque Map with project records.

Leak test associated exhaust ductwork.  Typical acceptance criteria will show no measurable leaks at 0.5 psig positive or negative.

Install fluid to appropriate levels to adjust glovebox to desired pressure levels as indicated by Magnehelic gauges.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall determine the pressure requirements and applicability of utilizing vacuum regulators or other devices to control inert atmosphere gloveboxes.  Refer to the American Glovebox Society, AGS-G001 and Figure 5.22 for a typical glovebox inert atmosphere with pressure control system.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Pressure Control Devices

Install pressure sensing equipment tied to vacuum breaker and alarm monitors to control inert gas supplies.

Install input and output gages as close to the pressure control device as possible.

Identify control devices in accordance with P&ID labelling.

Follow additional installation instructions in accordance with manufacture instructions.

Leak test associated exhaust ductwork.  Acceptance criteria will show no measurable leaks at 0.5 psig positive or negative.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall refer to the LANL Master Specifications, Division 26-Electrical for the relevant installation information listed in Section 2.2.G through 2.2.L.  The author shall also ensure electrical penetrations are hermetically sealed, and with electrical feedthroughs that are bench or factory leak tested with helium, to the minimum requirements (i.e., 1 X 10(-8) STD cc/sec.  Refer to Section 11 5311.16 – Hermetically-Sealed Glovebox Feedthroughs

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

electrical services

Connection to Circuits

Documentation

Hardwire all connections to circuits.

Identify conductors within specific electrical enclosures on Wiring Diagrams.

Labeling

Label electrical components.  Label electrical conductors with unique identification numbers.

Appropriately label conduits located behind gloveboxes, airlocks and other obstructions.

Power to equipment or electrical/instrumentation racks must meet national Electrical Code (NEC) requirements.

General Installation Guidelines

115 V power should be hard wired through conduit or raceways to a junction box mounted on the exterior of the glovebox, and directed through flexible insulated cable from the interior duplex outlet to the service.

Power (208 or 480 VAC) should be hard wired to junction boxes interior and exterior to the glovebox, with hermetic sealed and insulated pass-through.

Ensure there is a minimum 30-in of space available in front for access covers on junction boxes and electrical enclosures.

Ensure that access covers on junction boxes and electrical enclosures located between gloveboxes face the glovebox aisle.

Ensure that opposing electrical junction boxes for receptacles on glovebox service panels have a minimum of 4-in clearance between opposing cover plates.

Install low and moderate power feed-throughs on the top or side service panels of gloveboxes.

Installation Inside Gloveboxes

Install electrical connectors for systems operating in excess of 110 volts with a “twist-lock” feature.  Electrical plugs could come loose from electrical receptacles.

Installation Outside Gloveboxes

Completely enclose electrical raceways located underneath the glovebox.

Install NEMA 10 raceways for applications where there is a likelihood of acids or bases penetrating the interior of the raceway.

Completely enclose electrical raceways located above the glovebox.

Recess electrical racks a minimum of 6-in. back from the face of the glovebox.

Electrical raceways and conduit should not go up the front, operating side, or be located directly over a glovebox window or connector ring gasket.  Do not interfere with access to windows, vision into box, or activity, or interfere with action of gloves.

277/480-VAC Power

Follow facility specific color code.

120/208-VAC Power

Follow facility specific color code.

Lighting

Utilize circuits dedicated for wiring.

Coil up electrical cords near the plug-end of the cord and bind coil together with several nylon wire-ties.  This will increase access to the top of the glovebox for inspection, cleaning, and decontamination.

Install non-combustible lighting fixtures around hydrogenous shielding with “ring-shaped standoff” features.  Heat generated from high intensity lamps can melt the shielding material.  Replace single standoff pegs provided by the manufacturer of high intensity lighting with a loop or ring to prevent process personnel from injury.

Install non-combustible lighting fixtures for glovebox wells and intermediate floors.

Install non-combustible lighting fixtures with rubber standoff spacers.

Place non-combustible lighting fixtures with approximately 4-in. of clearance around the perimeter of the fixture.

Grounding

Properly ground glovebox support stands to the building ground grid, using heavy gauge braided copper cable, bolted (not clamped) to the steel structure box or support stand assembly.  Bond glovebox support stand to building ground electrode system with a #4 American Wire Gage (AWG) copper conductor.

Route ground cables around process equipment placed underneath glovebox.

Do not restrict maintenance access to ground cables.

instrumentation and controls

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall determine the pressure requirements and applicability of utilizing a Photohelic gauge to control inert atmosphere gloveboxes.  The system is typically used and supported by a pressure relief device (bubbler) in order to maintain prescribed pressure performance.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Refer to Section 11 5311.17 – Glovebox Instrumentation

Magnehelic / Photohelic Gauges

    - Magnehelic gauges provide localized indication of differential pressure between the glovebox and the laboratory work environment.  Mount on each side of the glovebox to provide a negative pressure potential for up to –2.0-in. water gauge.

    - Photohelic gauges control (2) solenoids on a bubbler bypass and inert gas supply system.  Gas feed switch is normally closed, while the pressure-reducing device (bubbler) is normally open.  Typical lower and upper negative pressure limit set points ranges are (0.3-0.5) and (0.5-0.7).  Install Gauges: As shown on installation drawings.

    - “Zero” Magnehelic and Photohelic or other pressure controllers and test for proper operation of makeup and exhaust solenoid valves in order to respond to changes in pressure.

Oxygen Monitors (Inert Atmosphere Gloveboxes Only)

O2 analyzer and moisture monitors may be incorporated into the purification system to indicate vapor-air content and maintain dilution of flammable mixtures to less than 25% of applicable lower combustible/explosive limits of materials processed or used in the process.

The Oxygen monitor may be an integral part of a Dri-Train unit necessary to maintain vapor-air content.  The unit can periodically or continuously draw atmospheric samples from glovebox interior through a pump to the external O2 analyzer.

Process samples of 1.0-ppm sensitivity are attainable with inert gas purge, purification system (Dri-Train) and O2 supply as appropriate.

These systems contain automatic and manual controls for adjusting or controlling the pressure in the glovebox (see Part A, “Magnehelic/ Photohelics”).

O2 monitors may be galvanic (depleting), or electrochemical type.

For fire protection limits, the oxygen threshold is about 5% for ordinary combustibles and ≤4% for pyrophoric metals.

The O2 monitor used to verify inert environments credited for fire prevention is required to alarm locally and at a constantly attended location.

Leak test piping to ensure system is free of leaks.

Reference AGS-G001 Section 5.6 – Atmosphere Systems, and Appendix A, Drawing AGS014, “Illustration of an Inert Atmosphere Glovebox with Oxygen Analyzing Equipment”).

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall determine and specify the radiation and criticality safety instrumentation requirements to address the issues below.  Analyses of glovebox and laboratory operations shall be conducted to identify the type, quantities, and placement of different instruments.  The author shall consider options for networking these instruments to provide increased flexibility and communication.  For criticality considerations, the author shall provide maximum dimensions and configurations of items containing fissile materials.

Analyses of glovebox and laboratory operations shall be conducted to identify the type, quantities, and placement of different instruments.  The design and analyses shall be reviewed by the Radiation Protection group in accordance with P 121, Radiation Protection.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

radiation and criticality safety

Personnel Contamination Monitors

Typically, a glovebox monitor is installed so that a hand probe is located between every other workstation (glove pairs).  Affix the probe to the glovebox to permit hands-free monitoring.  Mount the probe so that it does not interfere with normal glovebox operations.

Install hand and foot monitors at or near all exits from the laboratory to detect and prevent the spread of contamination beyond room boundaries.

Install hand and foot monitors on the “hot” side of step-off areas near the facility exit(s).

Hard-plumb supply of P-10 counting gas (90% argon, 10% methane) to the whole body contamination monitors.

Fixed Head Air Sampling (FAS)

Install a means to regulate the airflow through each FAS.  Often, a ball or needle valve is installed at each FAS location.  A flow meter is then connected in line with the FAS, and the valve is manually adjusted until the proper flow rate is established.

Optimally place FASs throughout the room to ensure that representative sampling of the airborne concentration is obtained.  Typically, one FAS is installed for every two workstations (glove pairs) mounted 6.5-7 feet off the floor.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall specify the make and model number of the CAM instrument.  A specification has been developed that identifies the requirements for a CAM used to monitor around operations with transuranic materials.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Continuous Air Monitors and Alarms

Continuous air monitors (CAMs) are used throughout glovebox areas Install a means to regulate the airflow through each CAM.  Often, a ball or needle valve is installed at each CAM location.  A flow meter is then connected in line with the CAM, and the valve is manually adjusted until the proper flow rate is established.  Some CAMs come with their own flow measurement system.

Install electrical power at each CAM location.  The facility engineering staff shall determine the specific power requirements based on the CAM selected for use.  Often, dedicated circuits are used for the CAM system.

Install a communications link at each CAM to facilitate remote monitoring.  This is typically a management decision weighing the need for remote/centralized notification, diagnostics, data trending and emergency response against the additional cost and maintenance.  Some newer CAMs are equipped to use the standardized RadNet communications protocol that allows networking the CAMs over a standard Ethernet local area network.

Area Radiation Monitors

Area Radiation Monitors (ARMs) should be installed in frequently occupied locations with the potential for unexpected increases in dose rates.  Also, ARMs should be installed in remote locations where there is a need for local indication of dose rates prior to personnel entry.

Install Central, remote indication of the dose rates measured by ARMs.  This remote monitoring will include an information transfer protocol such as RadNet to make all types of remote radiological monitoring consistent.  Archiving of dose rates measured by the ARMs should be included in the remote monitoring system.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The following section has been developed based upon procedures used at TA-55 PF-4 Facility.  Gloveboxes at LANL, however, regardless of the facility of installation, shall undergo a criticality safety review and criticality safety limits shall be posted on gloveboxes.

The author shall determine, develop and approve process accountability flow diagrams for transactions involving Material Accountability and Safeguards System (MASS).  The MASS at TA-55 is incapable of identifying approved material forms, material mixtures, container spacing, approved equipment, etc. that enter into the determination of criticality safety limits.  However, MASS listings are a useful administrative tool to assist nuclear material handlers in managing fissile material inventories within approved limits.  The author shall request addition of the glovebox to the facility’s MASS and a location name for the glovebox by formally submitting a request through SAFE-IP, Material Control and Accountability group, to SMS nuclear material management accountability group, to create the glovebox on MASS and receive a location name.

Procedures should be developed for approval of criticality limits (based on LANL Criticality Safety Group guidance, memo HSR-6-98-088 or successor) for two-tray carts, wagons, and flatbed carts.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Mass Accountability and Safeguards

For new or modified gloveboxes handling fissile materials, ensure criticality safety by performing a criticality review of the glovebox and glovebox operations.  Include the glovebox project supervisor, criticality safety officer, operations personnel, and Nuclear Criticality Safety group, in the criticality review and make recommendations to ensure criticality safety.

Incorporate the glovebox into the facility’s Material Accountability and Safeguards System (MASS).

Establish the following for gloveboxes and glovebox operations:

Isotopes and amounts of materials to be handled in the glovebox

Criticality safety limit/pass-through limits

Process description

Hazards

Diagram of glovebox work stations

Post workstation stations signs on gloveboxes that contain the following information:

Work station diagram, with criticality limits listed outside the diagram

LANL group number and issue date

MASS location identification

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Criticality tags are an operator aid used to represent, as accurately as possible, the total fissile material and form present in the glovebox workstation.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Ensure that tag (peg) boards are attached to the glovebox face and located as close as practical to the criticality signs on workstations.  Use only tag boards with CRITICALITY LIMITS at the top.  Use only one tag board for each workstation.

Install protective caps to prevent fire sprinkler water introduction into the glovebox, if applicable.

Radiation Shielding

Photon and neutron shield materials are used on and around gloveboxes to aid in maintaining radiation exposures as low as reasonably achievable.  The health physics staff shall be consulted to determine and specify shielding requirements.

Any permanent or temporary shielding materials shall be noncombustible.  Where noncombustible materials cannot be used, fire protection measures shall be provided as determined by the fire hazards evaluation.

The shielding type, configuration, thickness and material shall be specified based on the source term involved and the personnel exposure time.  Common materials for shielding of gloveboxes include stainless steel, lead, water, water extended polymer, and borated components.

Install shielding as close to the plutonium source as possible.

Install simple iron or lead shields placed over the gloveports to reduce the dose rates near the gloveboxes by an order of magnitude.

Tritium Monitors

Air Monitoring

Fixed ionization chamber instruments are the most widely used instruments for measuring gaseous forms of tritium in laboratory and process monitoring applications.  Portable ionization chamber instruments (i.e. tritium sniffers) are also used to control contamination and to supplement fixed instrument measurements.

For real-time air monitoring for tritium, the practical lower limits of sensitivity for these instruments range from 0.1 to 10 μCi/m3.  External background radiation can lower the sensitivity of these instruments.

Because tritium water vapor (HTO) is 10,000 to 25,000 times more toxic than elemental tritium (HT), it is often desirable to know the relative amounts of each species in the air.  Typically, such differential air monitoring is conducted with a single instrument that measures the total tritium concentration and then oxidizes the elemental tritium and repeats the measurement.

Ionization chambers are also used for process monitoring in hoods, exhaust systems, and gloveboxes.  The outputs are typically used to sound alarms, activate ventilation valves, or initiate some form of detritiation system.

Surface Monitoring

Tritium is difficult to measure directly because the low-energy beta radiation is readily absorbed in air and the in the window of the detector.  Normal frisking methods, such as the use of a pancake probe, will not detect tritium contamination.

For health and safety implications, an indication of loose, removable tritium contamination is more valuable than a measurement of total surface contamination.

Removable tritium contamination levels are routinely monitored by smears, which are wiped over a surface and then analyzed by liquid scintillation counting.  The smears are typically small round filter papers used dry or wet.  Wet smears are more efficient in removing tritium and the results are more reproducible, although the papers are usually more fragile when wet.  Removable tritium contamination measurements are only semi-quantitative, and reproducibility within a factor of 2 is considered satisfactory.

Ordinarily, an area of 100 cm2 of the surface is wiped with the smear paper and quickly placed in a vial with about 10 mL of liquid scintillation cocktail, or 1-2 mL of water with the cocktail added later.  The paper must be placed in liquid immediately after wiping because losses from evaporation can be considerable, especially if the paper is dry.

Personnel Monitoring

No reliable method of frisking personnel (skin or protective clothing) for tritium contamination is available.  Instead the best method used for evaluating personnel contamination is through bioassay (urinalysis).

Routine urine samples are collected at some predetermined frequency (typically bi-weekly) and counted for the tritium content.  This provides a very sensitive measurement of tritium in the body.

In addition to routine sampling, special tritium bioassays are collected whenever a potential exposure to tritium is suspected.

Criticality Alarm System

In most glovebox operations, the risk of an inadvertent criticality is very low; however, this risk cannot be eliminated.  It is often important to provide a means of alerting personnel of a criticality accident so they may promptly evacuate and limit their radiation exposure. The health physics and criticality safety staff shall conduct an evaluation to determine if criticality accident detection devices and alarm systems are needed.

The specification, installation, and use of a criticality alarm system (CAS) shall be established such that the total risk to personnel will be reduced.  Consideration shall be given to hazards that may result from false alarms.

The need for a criticality alarm system (CAS) shall be evaluated for all activities in which the inventory of fissionable material in individual, unrelated work areas exceeds 700 g of 235U, 520 g of 233U, 450 g of 239Pu, or 450 g of any combination of these three isotopes.

If the fissionable mass exceeds these limits and the probability of criticality is greater than 10-6 per year, a criticality alarm system shall be provided to cover occupied areas in which the expected dose exceeds 12 rad in free air.

The CAS should include criticality detection equipment and a personal evacuation alarm.

The alarm signal shall be for immediate evacuation purposes only and of sufficient volume and coverage to be heard in all areas that are to be evacuated.

The alarm signal shall activate promptly (i.e. within 0.5 sec) when the dose rate at the detectors equals or exceeds a value equivalent to 20 rad/min at 2 m from the reacting material.

The alarm signal shall be automatically actuated by an initiating event without requiring human action.

A visible or audible warning signal shall be provided at a normally occupied location to indicate system malfunction or loss of primary power.

The installation of the system shall be such as to resist earthquake damage.  The CAS should remain operational in the event of seismic shock equivalent to the site-specific design basis earthquake, or the equivalent value specified in the Uniform Building Code.

The location and spacing of the CAS detectors should be chosen to avoid the effect of shielding by massive equipment or materials.  Low-density materials of construction, such as wood framing, thin interior walls, hollow brick tiles, etc., may be disregarded.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The author shall edit the following section to incorporate requirements as defined in the facility authorization and design basis.  Elements form a general framework to be included in the specification and subsections can be added or deleted based on the specific installation.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Fire Safety

Combustible Loading

Maintain combustible loading within the GB, which is not stored within a noncombustible container to levels as low as possible. Limit non-containerized combustible loading to that required to support the day’s work

Maintain combustible loading surrounding the glovebox as low as possible.

Fire Suppression

Install Fire Suppression in gloveboxes, trolleys, drop boxes or hoods as required by the fire hazard evaluation (FHE); see NFPA 801 and AGS-G010.  Install a fire suppression system that is compatible with the glovebox operation/environment, criticality safety, combustible materials and material to be protected.  The fire suppression system shall be capable of notifying building occupants and the fire department of a system discharge.

Fire Screens

Install fire screens to protect ventilation Exhaust HEPA filtration. Firmly attach fire screens to the glovebox shell by a method appropriate to the materials of construction.  Refer to LANL drawing 26Y-202057 for details on HEPA filter housings and fire screens.

Install fire screens on zipper windows as indicated on LANL drawing 26Y-202006.

Fire Detection

Install fire detection as required by the glovebox or process FHE.  A fire detection system shall be selected that is compatible with the existing fire alarm control panel, the glovebox process and sufficient to address the fire risk and provide timely notification of a fire event.

Refer to Section 11 5311.12 – Glovebox Instrumentation for requirements pertaining to thermal detectors.

Install thermal detectors in the top of the glovebox at 8 ft spacing, starting no farther than 4 ft. from a glovebox wall, in a LANL-approved well as described in LANL Drawing 26Y-202010.

Where a tamper device is required, use an approved method to positively tie the installed detector to the glovebox such that the thermal detector may not be removed without evidence of the removal.

Test the thermal detection for alarm activation and active initiation of protective features as required by NFPA 72 and following an approved test procedure.

Egress and Maintenance Access

Install gloveboxes and other related equipment so that 36” egress aisle ways are maintained in accordance with International Building Code (IBC), International Fire Code (IFC). NFPA 101, Code for Safety to Life from Fire in Buildings and Structures (Life Safety Code).

Welding

Perform welding of carbon steel and stainless steel glovebox components in accordance with ESM Chapter 13 and Section 11 5311.10 – Glovebox Fabrication.

Provide welding documentation with project record documents, in accordance with ESM Chapter 13 and Section 11 5311.10, including Welding Procedure Specification, Welding Procedure Qualification Records, and Welder Qualification Test Records.

Perform weld examination in accordance with ESM Chapter 13 and Section 11 5311.10 – Glovebox Fabrication.

Welding performed inside a LANL facility shall be conducted in accordance with LANL Welding, Cutting, Other Spark- or Flame-Producing Operations (P101-26) requirements.

Painting

Paint carbon steel glovebox support stands in accordance with the requirements of 11 5311.10, following all installation fabrication and erection.

post-installation inspection and testing

Documentation and Reporting

Provide inspection and testing reports for all post installation examination, inspection and testing as indicated below.

Personnel Qualifications

Use only personnel trained to the same level as that required in Section 11 5311.10, Gloveboxes, for post installation inspection and testing.  Use only equipment calibrated per the requirements of Section 11 5311.10.

Visual Weld Examination

Confinement Welds: Visually inspect all confinement welds performed during installation on the glovebox and appurtenances in accordance with requirements defined in AWS D1.1.  Inspect welds from the inside of the confinement and from the outside of the confinement, where feasible.  Use the supplier provided weld map to locate seam welds that have been ground and polished. Note that there are often seam welds on the underside of the box.  Removal of the box from the stand is not required to inspect the covered weld from the outside unless other indications are present that make the integrity of this weld suspect.

Stand Welds:  Visually inspect welds performed during installation of the glovebox stand in accordance with requirements defined in AWS D1.1.  Inspect welds from both sides, where feasible.  Neither removal of the glovebox from the stand nor removal of legs from braces are required to inspect covered welds from the second side unless other indications are present that make the integrity of this weld suspect.

Repair welds not passing visual inspection using welding procedures defined in Section 11 5311.10.

Liquid Penetrant Examination

Confinement Welds: Liquid penetrant test all glovebox primary confinement welds performed during installation in accordance the liquid penetrant examination requirements identified in Section 11 5311.10.  Mark all primary confinement welds performed during installation on the weld map supplied with the glovebox developed during glovebox fabrication.

Repair welds not passing liquid penetrant examination in accordance with requirements set forth in Section 11 5311.10.

Surface Finish Inspection

Verify surface finish of glovebox in accordance with requirements defined on Drawing 26Y-202001.

Perform visual and tactile inspection of the entire glovebox exterior and spot check of accessible interior areas paying particular attention to any burrs or sharp edges that could inadvertently damage a glove.  Look especially for deep scratches caused by shipping or installation damage or by components shipped inside the glovebox breaking loose during transit.

Perform visual and tactile inspection of entire glovebox exterior and spot check of accessible interior areas paying particular attention to any areas of surface finish that may have decontamination concerns.

Perform surface finish inspection with a calibrated profilometer at several randomly selected areas on the interior and exterior of the glovebox.  Include any suspect areas found above.

Repair damage to the glovebox surface finish in accordance with the requirements set forth in Section 11 5311.10.

Inspect gloveports for any burrs or sharp edges that may have been caused during installation.

Fastener Inspection

Stand-to-Stand Connections: Visually inspect all stand-to-stand (legs, braces, and top) connections to verify that all split lock washers are in place and crushed and that there are no visible gaps greater than 1/16-inch between adjacent structural members.

Stand-to-Glovebox Connections: Visually inspect all stand-to-glovebox connections to verify that all fasteners are in place.  Verify by touch that no weld studs have been inadvertently broken by over tightening or shipping.

Stand-to-Floor Connections: Verify installation of anchor bolts in accordance with project documentation.

Window Clamp Strips: Visually inspect all bolted window clamp strip to glovebox connections to verify that all fasteners are in place.  Verify by touch that no weld studs have been inadvertently broken by over tightening or shipping.

Miscellaneous: Visually inspect all other threaded mechanical connections to verify that all fasteners are in place.  Verify by touch that no weld studs have been inadvertently broken by over tightening or shipping.  Pay particular attention to joints that affect confinement such as flanges between gloveboxes, filter housing attachments, and bolted service panels.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

The extent of piping examination, inspection and testing is defined in ASME B31.3.  The scope of examination, inspection and testing is dependent upon the service category of the piping system (Normal Service Category, Category D, Category M, and High-Pressure Service Category).  It is recommended that all piping systems transporting radioactive materials or have the potential of being contaminated with radioactive materials be categorized as Category M Services.  Category M services are defined as piping systems transporting hazardous materials and typically require more stringent leak integrity.  A helium leak test of piping systems can be substituted for the sensitive leak test requirements defined in ASME B31.3 provided that the helium leak rate is adequately defined.

(((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Piping Examination, Inspection, and Testing

All Piping Service Categories:

Perform piping examination, inspection and testing in accordance with the specific piping service category [Normal Service Category, Category D, Category M, or High-Pressure Service Category] requirements of ASME B31.3.

Follow ESM Ch 17 Pressure Safety, Section ADMIN-4, Inspection and Testing, for requirements and qualifications of personnel.

Comply with methods for performing inspection, examination and testing as outlined in ASME B31.3, including but not limited to visual weld examination, and leak testing.  Perform a pneumatic leak test in accordance with ASME B31.3 using design pressures for piping systems indicated above.  Submit procedures to LANL for inspection, examination and testing defined in ASME B31.3.

Use testing personnel qualified in accordance with ASME B31.3.  Use welding inspectors that have a current Certified Welding Inspector (CWI) certification in accordance with AWS QC-1 and ESM Chapter 13.

Generate test records for examination, inspection and testing in accordance with ASME B31.3 and submit the test records to LANL for review and approval.

Personnel performing Nondestructive Examination (NDE) or Nondestructive Testing (NDT), and the helium leak testing operations shall be qualified and certified as Level II, Level III or a Level I working under the direct supervision of a Level II or Level III in all applicable test methods, in accordance with the recommended practices of ASNT SNT-TC-1A or CP-189 and TA55-AP-025, Certification of Testing Personnel.

Category M Fluid Services

Develop and submit to LANL for review and approval procedures for performing examination, testing and inspection outlined in ASME B31.3 including, but not limited to, visual examination, visual weld examination, NDE, and leak testing.

Inspect, examine and test Category M service piping systems in accordance with the requirements set forth in ASME B31.3.  Perform pneumatic leak tests and a sensitive leak test using design pressures indicated above.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Select welds to be radiographed following welding by the supplier. Where radiography and ultrasonic inspection is not practical the use of in-process inspection requires planning in the field to have the inspectors present during fabrication instead of after.

((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((((

Perform random radiography of welds selected by LANL in accordance with ASME B31.3.  In addition to the requirements in ASME B31.3, perform in-process examination of 20% of welded joints using personnel other than those performing the production work. Where radiography is not practical, use ultrasonic or in-process inspection of welds in accordance with ASME B31.3

Use testing personnel qualified in accordance with ASME B31.3.  Use welding inspectors that have a current Certified Welding Inspector (CWI) certification in accordance with ESM Chapter 13 and AWS QC-1.

Generate test records for examination, inspection and testing in accordance with ASME B31.3 and submit the test records to LANL for review and approval.

Grind out and re-weld, using the original Welding Procedure Specification, welds not meeting the acceptance criteria for radiographic/ultrasonic examination as stated in ASME B31.3.  Re-examine 100% of repaired welds in accordance with ASME B31.3.  Do not repair welds more than twice before the section is abandoned and replaced.

Personnel performing Nondestructive Examination (NDE) or Nondestructive Testing (NDT), the helium leak testing operations shall be qualified and certified as Level II, Level III or a Level I working under the direct supervision of a Level II or Level III in all applicable test methods, in accordance with the recommended practices of ASNT SNT-TC-1A or CP-189 and TA55-AP-025, Certification of Testing Personnel.

Helium Leak Testing

Verify that all components of the glovebox system are installed prior to initiation of helium leak testing.

Isolate the glovebox from adjacent systems as required to contain the helium test gas. This may include blanking off the exhaust filter outlet, air inlets, and other openings.

Following installation and connection of the glovebox to existing facility systems, perform helium leak testing of the glovebox in accordance with the requirements in Section 11 5311.10.

The final helium leak test shall be performed after all GB utilities have been connected, all penetrations have been installed, and the GB installation in the operating facility is complete. The removal and reinstallation (including rework and repair) of a sealing component from the confinement boundary of the GB will negate the validity of any previously performed helium leak test and shall require a final helium leak test to be performed. The final helium leak test shall be performed prior to connection to the ventilation system.

Continuity Check

Following all wiring and labeling activities, perform electrical continuity testing.  Use a battery-powered buzzer or ohmmeter to determine that all power, control, instrumentation, and grounding, circuits are complete and properly identified.  Verify that the resistance between the main ground connection and the glovebox support stand and any point in the ground circuit is less than 0.1 ohms.  Verify instrument ground is separate and isolated from the power ground circuit and that no unintentional electrical grounds exist.  Prepare a log of circuits tested and submit the log to LANL for review and approval.

Checklist

Complete the Glovebox Certification Checklist provided in Attachment 1 prior to use of the glovebox.

END OF SECTION

******************************************************************************

Do not delete the following reference information.

******************************************************************************

FOR LANL USE ONLY

This project specification is based on LANL Master Specification 11 5311.12 Rev. 3 dated February 24, 2016.

| 1.0 Electrical Services |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 1.1 |  |  |  | Circuits are properly identified at penetrations |
| 1.2 |  |  |  | Circuits are properly identified at breaker panels |
| 1.3 |  |  |  | Voltage and wiring has been checked |
| 1.4 |  |  |  | Instrumentation receptacles and cables identified |
| 1.5 |  |  |  | Instrument pass-through are site or factory tested and properly installed |
| 1.6 |  |  |  | Photohelic and solenoid wiring has been checked |
| 1.7 |  |  |  | Glovebox frame properly grounded |
| 1.8 |  |  |  | Other |
| Electrical CSE | Signature/Date |

COMMENTS AND EXPLANATIONS:

__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 2.0 Fire Safety |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 2.1 |  |  |  | Fire detection is connected. Operability test must be performed in accordance with approved and applicable LANL Fire Alarm System acceptance testing procedures |
| 2.2 |  |  |  | Fire screens protect the HEPA filters |
| 2.3 |  |  |  | HEPA inlet filters have protective caps to protect against sprinkler discharge |
| 2.4 |  |  |  | Combustible loading is acceptable |
| 2.5 |  |  |  | Gloves are protected from hot equipment or operations |
| 2.6 |  |  |  | Equipment and cooling water circuits have over-temperature and loss of flow interlocks |
| 2.7 |  |  |  | Sprinkler protection or fire skirting installed beneath applicable glovebox |
| 2.8 |  |  |  | Furnaces have over temperature shutoffs that require manual intervention to restart the power supply. |
| 2.9 |  |  |  | Other |
| Fire Protection CSE | Signature/Date |

COMMENTS AND EXPLANATIONS:

________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 3.0 Seismic Safety |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 3.1 |  |  |  | Glovebox support properly attached to the floor |
| 3.2 |  |  |  | Glovebox properly attached to the support |
| 3.3 |  |  |  | Glovebox support members properly attached |
| 3.4 |  |  |  | Check stand configuration against drawings |
| 3.5 |  |  |  | Heavy objects on top of or inside glovebox properly secured |
| 3.6 |  |  |  | Other |
| Glovebox Support Stands CSE | Signature/Date |

COMMENTS AND EXPLANATIONS:

__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 4.0 Glovebox Integrity |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 4.1 |  |  |  | Glovebox interior meets surface finish and flatness requirements as defined on Drawing 26Y-202001. Verify flatness around openings prior to installation and following any welding performed on glovebox shell during installation. |
| 4.2 |  |  |  | Fabricator certified that glovebox is helium leak-tight (new installations only) |
| 4.3 |  |  |  | Helium leak test performed after all GB utilities connected, penetrations installed, and glovebox installation is complete (new or cold installations only). The removal and reinstallation (including rework and repair) of a sealing component from the confinement boundary of the GB will negate the validity of any previously performed helium leak test and shall require a final helium leak test to be performed. |
| 4.4 |  |  |  | Service panels, filter hat and flanged penetrations have approved gaskets |
| 4.5 |  |  |  | Window frame bolts have been torqued (25 ( 5 in-lb) and gasket surfaces appear smooth. Supply Torque Map to document fastener torques. |
| 4.6 |  |  |  | Zipper window gaskets have proper zippers |
| 4.7 |  |  |  | Unused penetrations are plugged, preferably inside and outside |
| 4.8 |  |  |  | Gloves are properly attached to glove rings |
| 4.9 |  |  |  | HEPA filters to Zone 1 exhaust have been properly installed and successfully leak tested |
| 4.10 |  |  |  | Hoods have the required face velocity (HSR certified) |
| 4.11 |  |  |  | Glovebox pressure gauges (Magnehelic or Photohelic) are active and respond to changes in pressure |
| 4.12 |  |  |  | Glovebox pressure gauges (Magnehelic and Photohelic) have been “zeroed” and accurately reflect pressure in the glovebox |
| 4.13 |  |  |  | Other |
| Glovebox CSE | Signature/Date |

COMMENTS AND EXPLANATIONS:

__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 5.0 Mechanical Services |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 5.1 |  |  |  | All gas and liquid services are labeled |
| 5.2 |  |  |  | Industrial water piped into glovebox has “dead-man” valves to prevent flooding |
| 5.3 |  |  |  | Gas service lines to gloveboxes are 3/8-in. or smaller except with special approval (Note: ¼-in. or smaller for helium) |
| 5.4 |  |  |  | No positive pressure cooling water, except for limited volume systems, is piped into gloveboxes |
| 5.5 |  |  |  | Positive pressure chilled circulating water is used to cool equipment exterior to the glovebox |
| 5.6 |  |  |  | Chilled or heated piping is properly insulated |
| 5.7 |  |  |  | House and wet vacuum lines are trapped to prevent flooding of lines or gloveboxes |
| 5.8 |  |  |  | All gas supplies entering gloveboxes have pressure regulating valves, check valves, gauges near work stations, properly sized orifices in services above 25 psig, and high flow, in-line filters |
| 5.9 |  |  |  | All pressure regulating valves (PRV) for gas services are set for 25 psig maximum (except for instrument air, or by special approval) |
| 5.10 |  |  |  | At least one gauge is provided downstream of all PRVs to verify indications of the PRV gauges |
| 5.11 |  |  |  | Glovebox service lines have been traced to facility service lines to verify labeling |
| 5.12 |  |  |  | Flanges of wet vacuum and liquid-carrying lines are plastic-wrapped |
| 5.13 |  |  |  | Supply lines for air-driven stirrers, vibrators, pumps, etc., inside gloveboxes have solenoid valves interlocked to glovebox pressure |
| 5.14 |  |  |  | Air-operated cylinders have PRVs and flow-regulating valves |
| 5.15 |  |  |  | Air-operated cylinders with shafts penetrating contaminated gloveboxes have HEPA filters and exhaust to Zone 1 ductwork |
| 5.16 |  |  |  | Vacuum pumps have metal piping to gloveboxes, have an intake HEPA filter (in accordance with Section 23 4133) between the vacuum pump and glovebox, and exhaust through metal piping to the Zone 1 system |
| 5.17 |  |  |  | HEPA filter housings (meeting the requirements of Section 23 4133) appear in good condition |
| 5.18 |  |  |  | Pressurized piping, vessels, columns, etc., have been evaluated and leak tested for operating pressures |
| 5.19 |  |  |  | Other |
| Glovebox CSE or Mechanical Engineer | Signature/Date |

COMMENTS AND EXPLANATIONS:

__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 6.0 Special Items for Inert Atmosphere Gloveboxes |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 6.1 |  |  |  | The damper in the Zone 1 exhaust line is locked in the full open position |
| 6.2 |  |  |  | Pressure relief devices (bubblers) are filled with approved fluid to the proper level |
| 6.3 |  |  |  | The glovebox, airlock, and connecting piping have been helium leak checked after installation (new and cold installations only) |
| 6.4 |  |  |  | Recirculating purifiers (Dri-Trains) have HEPA filters, PRVs and flow meters (where required) in gas supply lines |
| 6.5 |  |  |  | Dri-Train recirculating lines are hard soldered and equipped with high vacuum shut-off valves and inlet and outlet HEPA filters inside the glovebox |
| 6.6 |  |  |  | Dri-Train vacuum pumps have approved filters and exhaust into the Zone 1 system |
| 6.7 |  |  |  | Dri-Trains are of the type approved for PF-4 service |
| 6.8 |  |  |  | Other |
| Dri-Train System Engineer | Signature/Date |

COMMENTS AND EXPLANATIONS:

______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 7.0 Radiation and Criticality Safety |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 7.1 |  |  |  | Criticality assessment has been performed and approved criticality limits have been posted and pegboards are in place |
| 7.2 |  |  |  | Hand monitors are in place near work stations |
| 7.3 |  |  |  | Fixed head air samplers are properly mounted on the gloveboxes |
| 7.4 |  |  |  | CAM air sample tubes are in place at hood faces |
| 7.5 |  |  |  | Hand and foot monitors are available in the working area |
| 7.6 |  |  |  | LANMAS locations and glovebox numbers are posted |
| 7.7 |  |  |  | Radiation shielding is in place |
| 7.8 |  |  |  | Other |
| Glovebox CSE or Radiation Protection | Signature/Date |

COMMENTS AND EXPLANATIONS:

________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 8.0 Glovebox Glove Sharps (Ref. Lessons Learned # 2009-26) |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 8.1 |  |  |  | Area around glovebox gloves is clear of pinch points or unplanned contact with equipment |
| 8.2 |  |  |  | Equipment inside glovebox free from sharp corners, sharp edges, and abrasive surfaces |
| 8.3 |  |  |  | Threaded fasteners inside glovebox do not have more than 1 to 2 exposed threads |
| 8.4 |  |  |  | Coated or lined gloveboxes are free from sharp nicks |
| 8.5 |  |  |  | Other |
| Sharps Committee Representative | Signature/Date |

COMMENTS AND EXPLANATIONS:

__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| 9.0 Summary Certification Finding |
| --- |
| Item. | Initial Finding | Comment Y/N | Final Finding | Glovebox Requirement |
| 9.1 |  |  |  | New glovebox installation |
| 9.2 |  |  |  | Modified glovebox installation |

COMMENTS AND EXPLANATIONS:

____________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

-INSTRUCTIONS & ABBREVIATIONS-

Certification that gloveboxes, hoods, equipment, and mechanical and electrical services meet LANL TA-55 standards will be done by qualified personnel.

Inspections may be performed by individuals alone or by teams of qualified individuals.  Inspections should be performed during appropriate phases of construction to avoid costly changes after construction is completed.

Findings may be:

- “Conforming” (C),

- “Not Applicable” (NA),

- “Nonconforming - OK” (NC-OK),

- “Nonconforming - Must Correct” (NC-MC), or

- “Pending certification after connection to Zone 1” (P)

All findings, except “C”, must be addressed in the “Comments and Explanations”.

| Connection to Zone 1 (not for normal operations) is authorized: |
| --- |
| Glovebox CSE | Signature/Date |

| Glovebox meets facility requirements for normal operations |
| --- |
| Glovebox CSE | Signature/Date |